## Basic Physics
question | question a1 | a2 | a3 | a4 | a5 | answer of a1 | explanation of a1 | answer of a2 | explanation of a2 | answer of a3 | explanation of a3 | answer of a4 | explanation of a4 | answer of a5 | explanation of a5
Regarding fundamental MRI principles: | Does MRI primarily utilize the magnetic properties of electrons in hydrogen atoms? | In the absence of an external magnetic field (B₀), do proton spins align to create a strong net magnetization vector? | Is the net magnetization vector (NMV) oriented parallel to the main magnetic field (B₀) at equilibrium? | Are tissues with a high concentration of free water molecules poor sources of MRI signal? | Is the primary magnetic field (B₀) typically oriented perpendicular to the long axis of the patient in a cylindrical scanner? | FALSE | MRI utilizes the magnetic properties of atomic nuclei with an odd number of protons or neutrons, primarily the single proton in hydrogen atoms. | FALSE | In the absence of an external field, proton spins are randomly oriented, resulting in a net magnetization of zero. | TRUE | At equilibrium, the spins align with the B₀ field, creating a net magnetization vector (NMV) that is parallel to the direction of B₀. | FALSE | Tissues with a high density of mobile protons, such as those with high water content, are excellent sources of signal. | FALSE | In a standard cylindrical (solenoidal) scanner, the main B₀ field is oriented parallel to the long axis of the patient (the z-axis).
Concerning the Larmor frequency: | Is the Larmor frequency directly proportional to the strength of the gradient magnetic fields? | Is it defined as the precessional frequency of a nucleus in an external magnetic field? | For a given nucleus, does the Larmor frequency decrease as the main magnetic field (B₀) strength increases? | Is the gyromagnetic ratio a variable that changes based on tissue type for a given nucleus? | Are radiofrequency (RF) pulses applied at a frequency significantly different from the Larmor frequency to excite spins? | FALSE | The Larmor frequency is directly proportional to the main magnetic field strength (B₀), not the much weaker, superimposed gradient fields. | TRUE | This is the fundamental definition. The Larmor equation (ω₀ = γB₀) states that the precessional frequency (ω₀) is the product of the gyromagnetic ratio (γ) and B₀. | FALSE | The Larmor frequency is directly proportional to B₀. As the magnetic field strength increases, the precessional frequency increases. | FALSE | The gyromagnetic ratio is a physical constant for a specific nucleus (e.g., hydrogen-1) and does not change based on its chemical environment or tissue type. | FALSE | To achieve resonance and excite the spins, the RF pulse must be transmitted at precisely the Larmor frequency of the target nuclei.
On the subject of T1 relaxation: | Does T1 relaxation refer to the recovery of transverse magnetization after an RF pulse? | Is T1 relaxation also known as spin-spin relaxation? | Do tissues with a long T1 value, such as fat, recover their longitudinal magnetization quickly? | Is T1 relaxation primarily caused by the exchange of energy between excited spins and the surrounding molecular lattice? | Is the T1 time constant defined as the time it takes for 100% of the longitudinal magnetization to recover? | FALSE | T1 relaxation refers to the recovery of longitudinal magnetization along the z-axis. T2 decay concerns the loss of transverse magnetization. | FALSE | T1 relaxation is known as spin-lattice relaxation. T2 relaxation is known as spin-spin relaxation. | FALSE | Fat has a short T1 and recovers quickly. Tissues with a long T1, like cerebrospinal fluid (CSF), recover slowly. | TRUE | This is the definition of spin-lattice relaxation. Protons release energy to their surrounding molecular environment (the "lattice") to realign with the B₀ field. | FALSE | The T1 time constant is the time required for the longitudinal magnetization to recover to approximately 63% of its final equilibrium value.
With respect to T2 and T2 decay:* | Is T2 decay, or spin-spin relaxation, typically a much slower process than T1 relaxation in biological tissues? | Does T2 decay represent the loss of phase coherence among spins in the transverse plane? | Is T2* decay caused solely by molecular spin-spin interactions, excluding external field inhomogeneities? | Does a spin-echo sequence correct for signal loss due to both local field inhomogeneities and true spin-spin interactions? | Do tissues with a short T2, such as dense ligaments, appear bright on T2-weighted images? | FALSE | In biological tissues, T2 decay is always a much faster process than T1 relaxation (i.e., the T2 time constant is significantly shorter than the T1 time constant). | TRUE | T2 decay is the dephasing of spins in the transverse (x-y) plane due to their magnetic interactions with each other, leading to signal loss. | FALSE | T2* decay is caused by the combination of true T2 spin-spin interactions and dephasing from local magnetic field inhomogeneities. | FALSE | A spin-echo's 180° pulse refocuses signal loss from static field inhomogeneities but cannot correct for the irreversible signal loss from true T2 spin-spin interactions. | FALSE | Tissues with a short T2 lose signal rapidly and therefore appear dark on T2-weighted images. Tissues with a long T2 (e.g., fluid) remain bright.
Regarding MRI pulse sequences and parameters: | In a conventional spin-echo sequence, is the 180° pulse applied before the 90° excitation pulse? | Is the Repetition Time (TR) the time interval between the application of two consecutive 90° pulses in a sequence? | Is the Echo Time (TE) the time from the 180° refocusing pulse to the peak of the echo? | To achieve T1-weighting, should a long TR and a long TE be used? | Is Proton Density (PD) weighting achieved by minimizing both T1 and T2 effects, typically with a long TR and short TE? | FALSE | In a standard spin-echo sequence, the 90° excitation pulse is applied first to create transverse magnetization, followed by a 180° pulse to refocus the spins. | TRUE | This is the definition of TR. It is a key parameter that determines the amount of T1 relaxation that occurs between excitations and thus controls T1-weighting. | FALSE | The Echo Time (TE) is the time measured from the center of the initial 90° excitation pulse to the center of the resulting echo. | FALSE | T1-weighting requires a short TR (to emphasize differences in T1 recovery) and a short TE (to minimize T2 decay effects). | TRUE | A long TR allows most tissues to fully recover their longitudinal magnetization (minimizing T1 effects), and a short TE minimizes signal loss from T2 decay, thus isolating contrast based on proton density.
On the principles of spatial encoding: | Is the slice-selection gradient applied simultaneously with the 180° refocusing pulse in a spin-echo sequence? | Is the frequency-encoding gradient, also known as the readout gradient, turned on during the data acquisition window? | Is the phase-encoding gradient applied with a constant amplitude for each line of k-space? | Does the phase-encoding gradient cause spins to precess at different frequencies depending on their spatial location? | Is spatial encoding achieved without the use of magnetic field gradients in modern MRI systems? | TRUE | A slice-selection gradient must be applied during both the 90° pulse (to excite the slice) and the 180° pulse (to ensure only spins within that same slice are refocused). | TRUE | The frequency-encoding gradient is active while the echo is being received, causing spins at different positions along its axis to precess at different frequencies, which encodes their location. | FALSE | The phase-encoding gradient is applied with a different amplitude (stepped from positive to negative) for each repetition (each TR), imparting a unique phase shift to each line of the image. | FALSE | The phase-encoding gradient imparts a brief, position-dependent phase shift. It is the frequency-encoding gradient that causes spins to precess at different frequencies. | FALSE | Magnetic field gradients are the fundamental tool used to create spatial variation in the magnetic field, which is essential for encoding the location of the MR signal.
Pertaining to k-space: | Does each point in k-space correspond directly to a single pixel in the final MR image? | Is the center of k-space filled by data acquired with high-amplitude phase-encoding gradients? | Does the center of k-space primarily contain information about the image's contrast and overall brightness? | Do the peripheral regions of k-space contribute mainly to the signal-to-noise ratio of the image? | Is k-space traversed in a random pattern to reduce imaging time in standard acquisitions? | FALSE | K-space contains spatial frequency data. The entire k-space must be processed via a Fourier Transform to reconstruct the final image; there is no direct point-to-pixel correlation. | FALSE | The center of k-space, which represents low spatial frequencies, is filled using low or zero amplitude phase-encoding gradients. High-amplitude gradients fill the periphery. | TRUE | The low spatial frequencies located in the center of k-space encode the broad shapes and signal intensities, which determine the fundamental contrast of the image. | FALSE | The periphery of k-space contains high spatial frequencies, which are responsible for encoding the fine details and sharp edges in the image. | FALSE | In standard Cartesian imaging, k-space is filled in a sequential, linear order (line by line). Other non-linear or radial patterns exist for advanced techniques.
Examining common MRI artifacts: | Is aliasing, or wrap-around artifact, caused by setting a field of view (FOV) that is too large for the anatomy being imaged? | Does chemical shift artifact result in a misregistration of signal between fat and water, typically in the frequency-encoding direction? | Are motion artifacts typically seen as distinct, sharp ghosting patterns in the slice-selection direction? | Can magnetic susceptibility artifacts, caused by metal, be reduced by using gradient-echo sequences instead of spin-echo? | Is zipper artifact typically caused by patient motion during the scan? | FALSE | Aliasing artifact is caused when the Field of View (FOV) is smaller than the body part being scanned, causing anatomy outside the FOV to be mapped onto the opposite side of the image. | TRUE | The slight difference in Larmor frequency between fat and water protons causes a spatial shift artifact along the frequency-encoding axis. | FALSE | Because the phase-encoding steps are applied over time, periodic motion results in ghosting artifacts that propagate along the phase-encoding direction. | FALSE | Gradient-echo (GRE) sequences lack a 180° refocusing pulse and are therefore more sensitive to the local field distortions that cause magnetic susceptibility artifacts. | FALSE | Zipper artifact appears as a line of noise across the image and is caused by extrinsic radiofrequency (RF) noise leaking into the scanner room from an external source.
Regarding advanced imaging techniques: | Do gradient-echo (GRE) sequences use a 180° RF pulse to refocus spins and form an echo? | Are GRE sequences generally slower than conventional spin-echo sequences for acquiring a single slice? | Is Inversion Recovery (IR) a technique used to null the signal from a specific tissue by applying an initial 180° pulse? | Does Diffusion-Weighted Imaging (DWI) measure the T2 decay of water to create image contrast? | Is Magnetic Resonance Angiography (MRA) unable to visualize blood flow without the use of an intravenous contrast agent? | FALSE | GRE sequences use a bipolar gradient reversal to rephase the spins and form an echo, not a 180° RF pulse. This is a key difference from spin-echo sequences. | FALSE | GRE sequences are significantly faster because they can use smaller flip angles and do not require the time needed for a 180° pulse, allowing for much shorter TRs. | TRUE | By applying a 180° pulse and waiting a specific Inversion Time (TI), the magnetization of a target tissue can be nulled as it passes through zero, suppressing its signal. | FALSE | DWI uses strong gradients to measure the magnitude of random thermal motion of water molecules (diffusion). It is sensitive to motion, not just T2 decay. | FALSE | Non-contrast MRA techniques, such as Time-of-Flight (TOF) and Phase Contrast (PC), can visualize blood flow by exploiting the signal differences between flowing and stationary spins.
On the topic of MRI hardware components: | Are gradient coils responsible for generating the main static magnetic field (B₀)? | Is the RF coil used for both transmitting the excitation pulse and receiving the MR signal? | Are shim coils used to degrade the homogeneity of the magnetic field for spatial encoding? | Is the primary function of the Faraday cage to protect the patient from acoustic noise? | Are superconducting magnets cooled to near absolute zero primarily to reduce the energy consumption of the scanner? | FALSE | The main B₀ field is generated by the large superconducting magnet. Gradient coils produce small, rapid, linear variations in this field for spatial encoding. | TRUE | Many RF coils are transceivers, meaning they are capable of both transmitting the RF pulse and receiving the resulting signal. Some systems use separate transmit and receive coils. | FALSE | Shim coils are used to improve the homogeneity (uniformity) of the main B₀ field by applying small, static corrections to counteract imperfections. | FALSE | The Faraday cage is a copper-shielded room designed to block external radiofrequency noise from entering and corrupting the weak MR signal, thus protecting image quality. | TRUE | By cooling the coils to a superconducting state, electrical resistance drops to nearly zero. This allows a massive current to be maintained with minimal power to generate the strong B₀ field.
Concerning MRI safety considerations: | Is the "missile effect" a risk associated with the radiofrequency (RF) heating of tissue? | Is Specific Absorption Rate (SAR) a measure of the RF energy absorbed by the body per unit mass? | Are patients with non-ferromagnetic metallic implants, such as titanium, always contraindicated for MRI scans? | Can the rapidly switching gradient fields cause uncomfortable heating of the patient's skin? | Is Zone IV the most restricted area, which includes the scanner room itself? | FALSE | The missile effect is the risk of ferromagnetic objects being pulled into the scanner by the powerful static magnetic field (B₀). RF heating is a separate concern. | TRUE | SAR is the FDA-regulated measure of RF energy deposition in the body, expressed in watts per kilogram (W/kg). It is monitored to prevent excessive tissue heating. | FALSE | Many non-ferromagnetic implants, particularly those made of titanium, are considered MR-conditional or MR-safe. However, every device must be individually verified for safety. | FALSE | Rapidly switching gradients can induce electrical currents that cause peripheral nerve stimulation (twitching or tingling). Significant heating is associated with the RF field (SAR). | TRUE | According to the American College of Radiology (ACR) safety guidelines, Zone IV is the magnet room itself and is under the direct control of MRI technologists, requiring the strictest access control.
With respect to MRI contrast agents: | Do gadolinium-based contrast agents work by significantly increasing the T2 relaxation time of adjacent water protons? | Is gadolinium in its free ionic form (Gd³⁺) considered safe for direct administration to patients? | Do these agents primarily enhance tissues by shortening their T1 relaxation time, making them appear brighter on T1-weighted images? | Is Nephrogenic Systemic Fibrosis (NSF) a known risk associated with these agents primarily in patients with normal renal function? | Are MRI contrast agents universally required for all diagnostic MRI examinations? | FALSE | Gadolinium is a paramagnetic substance that creates local magnetic fields, which dramatically shortens (decreases) both the T1 and T2 relaxation times of nearby protons. | FALSE | The free gadolinium ion is highly toxic. For clinical use, it must be bound to a large organic molecule called a chelate to render it non-toxic. | TRUE | The dominant effect used for clinical enhancement is T1 shortening. Tissues that accumulate the agent recover their longitudinal magnetization faster and appear hyperintense (bright) on T1-weighted scans. | FALSE | NSF is a rare but severe fibrosing disease that has been linked to gadolinium-based contrast agents almost exclusively in patients with severe, pre-existing kidney disease. | FALSE | A large number of MRI studies are performed without contrast. Contrast is used for specific indications, such as evaluating tumors, infection, or vascular abnormalities.
Regarding the fundamental principles of atomic structure: | Is the atomic number defined as the sum of an atom's protons and electrons? | Are electrons distributed in discrete energy shells surrounding the atomic nucleus? | Within a specific element, do all electrons occupy a single, uniform energy level? | Is the physical size of a single nucleon less than that of the nucleus it constitutes? | Does the atomic nucleus possess an overall positive electrical charge? | FALSE | The atomic number (Z) is defined by the number of protons only. In a neutral atom, the electron count equals the proton count, but this does not define the atomic number. | TRUE | According to the Bohr atomic model, electrons exist in quantized energy levels, or shells (K, L, M, etc.), at specific distances from the nucleus. | FALSE | Electrons within the same atom occupy different shells, each with a distinct binding energy. A K-shell electron has a much higher binding energy than an L-shell electron. | TRUE | A nucleon is a single particle (proton or neutron) within the nucleus. For any atom other than Hydrogen-1, the nucleus contains multiple nucleons and is therefore larger. | TRUE | The nucleus is composed of positively charged protons and uncharged neutrons, resulting in a net positive charge.
Concerning the properties of elements pertinent to radiological applications: | Is tungsten (W) characterized by an atomic number of 82? | Is the K-shell binding energy (K-edge) for barium approximately 37 keV? | Is the K-shell binding energy (K-edge) for calcium (Ca) approximately 20 keV? | Does Technetium-99m ($^{99m}$Tc) exhibit a principal gamma emission energy of approximately 140 keV? | Does iodine possess an atomic number of 53? | FALSE | Tungsten (W) has an atomic number of 74. Lead (Pb) has an atomic number of 82. | TRUE | Barium's K-edge of ~37 keV is well within the diagnostic X-ray energy range, leading to a high probability of photoelectric absorption, which makes it an excellent contrast agent. | FALSE | The K-edge of calcium is much lower, at approximately 4 keV. | TRUE | The monoenergetic 140 keV gamma photon emission of $^{99m}$Tc is ideal for imaging with modern gamma cameras and is the most common radionuclide in nuclear medicine. | TRUE | Iodine has an atomic number of 53. Its K-edge of ~33 keV makes it an effective iodinated contrast agent for CT and angiography.
Regarding the characteristics of electromagnetic radiation: | Are radio waves classified as a form of electromagnetic radiation? | Is the propagation speed of radio waves in a vacuum 3×108 m/s? | Does the fundamental distinction between X-rays and gamma rays lie solely in their mechanism of production? | Is the relationship between frequency and wavelength one of direct proportionality? | Does the propagation of radio waves occur in straight, diverging paths? | TRUE | The electromagnetic spectrum includes radio waves, microwaves, infrared, visible light, ultraviolet, X-rays, and gamma rays. | TRUE | All forms of electromagnetic radiation travel at the speed of light (c≈3×108 m/s) in a vacuum. | TRUE | This is the defining difference. X-rays are extranuclear in origin (from electron interactions), while gamma rays are intranuclear (from radioactive decay). | FALSE | Frequency and wavelength are inversely proportional (c=fλ). As frequency increases, wavelength decreases. | TRUE | Like all forms of electromagnetic radiation, radio waves propagate from a source in straight lines that spread out (diverge) with distance.
With respect to the process of X-ray generation: | Does Bremsstrahlung describe the increase in kinetic energy of electrons during acceleration toward the anode target? | Does the phenomenon of thermionic emission take place at the target material? | Are Bremsstrahlung photons generated at the cathode assembly of an X-ray tube? | Does the generation of characteristic radiation occur at the anode of an X-ray tube? | Is approximately half of the kinetic energy from incident electrons converted into diagnostically useful X-ray photons? | FALSE | Bremsstrahlung ("braking radiation") describes the loss of kinetic energy by an electron as it is decelerated by the nucleus of a target atom, with the lost energy converted into an X-ray photon. | FALSE | Thermionic emission, the release of electrons due to heat, occurs at the filament of the cathode. | FALSE | Bremsstrahlung X-rays are generated at the target on the anode when high-speed electrons from the cathode strike it. | TRUE | Characteristic radiation is produced at the anode when a high-speed electron ejects an inner-shell electron from a target atom, causing a cascade of electrons that releases a photon. | FALSE | X-ray production is extremely inefficient. More than 99% of the electron's kinetic energy is converted into heat, with less than 1% forming X-rays.
On the subject of X-ray beam attenuation: | Is attenuation defined as the reduction of photon fluence in a beam due to absorption and scattering events? | For a monoenergetic beam, does the half-value layer (HVL) increase as a function of penetration depth? | Do digital radiography systems typically employ monoenergetic X-ray beams? | Is there an inverse proportional relationship between the linear attenuation coefficient and the half-value layer (HVL)? | Is an X-ray beam's photon count reduced to zero after passing through five half-value layers of a material? | TRUE | This is the correct definition. Attenuation is the removal of photons from the primary beam by either absorption (e.g., photoelectric effect) or scattering (e.g., Compton scatter). | FALSE | For a monoenergetic beam, the probability of interaction per unit length is constant. Therefore, the HVL is also constant, regardless of depth. | FALSE | Digital radiography systems use standard X-ray tubes, which produce a polyenergetic (or heterogeneous) spectrum of X-ray energies. | TRUE | This is a fundamental relationship, defined by the equation HVL = 0.693 / µ, where µ is the linear attenuation coefficient. | FALSE | Attenuation is an exponential process. Each HVL reduces the intensity by 50%, so the intensity approaches zero asymptotically but mathematically never reaches it.
Further considerations on X-ray attenuation: | In a clinical setting, are the first and second half-value layers of an X-ray beam equivalent? | Does a material's mass attenuation coefficient exhibit an inverse relationship with its atomic number? | Do elastic (coherent) scattering interactions have a significant impact on radiographic image quality? | Does the photoelectric effect result in the total absorption of the incident X-ray photon's energy? | Is the photoelectric effect the dominant interaction mechanism when diagnostic X-rays traverse high-Z contrast agents like barium? | FALSE | Clinical X-ray beams are polyenergetic. Lower-energy photons are removed first (beam hardening), so the average energy of the beam increases with depth. Thus, the second HVL is greater than the first. | FALSE | The mass attenuation coefficient generally increases with atomic number (Z), particularly for the photoelectric effect, which is proportional to Z³. | FALSE | Coherent (Rayleigh) scattering occurs at low energies and results in a small change in photon direction with no energy loss. It contributes minimally to image fog in the diagnostic range. | TRUE | In a photoelectric interaction, the incident photon is completely absorbed, transferring all its energy to an inner-shell electron, which is then ejected from the atom. | TRUE | The high atomic number (Z) of contrast agents like barium (Z=56) and iodine (Z=53) causes the photoelectric effect probability (proportional to Z³) to be very high, leading to excellent contrast.
Concerning the properties of a clinical X-ray beam: | Would the removal of characteristic photons from an X-ray spectrum result in an increase in the total beam intensity? | Would increasing the distance from a point source from 1 meter to 4 meters cause the beam intensity to decrease by a factor of 16? | Is the inverse square law predicated on the assumption of a non-attenuating medium? | Is the average photon energy of a polyenergetic X-ray beam approximately 50% of the peak kilovoltage (kVp)? | Is a clinically produced X-ray beam characterized by a constant half-value layer (HVL)? | FALSE | Beam intensity is a measure of the total energy flux. Removing any photons from the beam, including characteristic photons, would decrease the overall intensity. | TRUE | The inverse square law states that intensity is inversely proportional to the square of the distance (I∝1/d2). Increasing the distance by a factor of 4 decreases the intensity by a factor of 4², or 16. | TRUE | The law describes the geometric divergence of radiation from a point source and does not account for the absorption or scattering of photons by the intervening medium (e.g., air). | FALSE | Due to the shape of the Bremsstrahlung spectrum, the average energy of a filtered X-ray beam is typically about one-third (1/3) of the peak kilovoltage (kVp). | FALSE | Because clinical beams are polyenergetic, they undergo "beam hardening" as they pass through matter. This increases the average energy of the beam, which in turn increases the HVL with depth.
Regarding the principles of the Compton effect: | Is the Compton effect the mechanism describing the scattering of photons within biological tissue? | Does the interaction involve the ejection of a photoelectron following a collision with an incident photon? | Are the "free" electrons involved in Compton scattering typically outer-shell or valence electrons? | Is the probability of a Compton interaction proportional to the cube of the absorber's atomic number (Z³)? | For the Compton effect, is the mass attenuation coefficient of bone double that of soft tissue? | TRUE | Compton scattering is the dominant interaction in soft tissue for most of the diagnostic X-ray energy range, and it is the primary source of scatter radiation. | FALSE | Compton scattering results in a Compton electron (or recoil electron). The term "photoelectron" is reserved for the electron ejected during the photoelectric effect. | TRUE | The interaction occurs with loosely bound, outer-shell electrons whose binding energy is negligible compared to the incident photon energy, effectively treating them as "free." | FALSE | The Z³ dependence is characteristic of the photoelectric effect. Compton scattering probability is nearly independent of Z and is instead proportional to the material's electron density. | FALSE | The mass attenuation coefficient for Compton scatter depends on electron density per gram, which is very similar for bone and soft tissue. Therefore, their coefficients are nearly equal.
Examining the photoelectric effect: | Does the maximum probability for the photoelectric effect occur when the incident photon energy is marginally below the electron's binding energy? | Can the emission of characteristic radiation be a subsequent event following a photoelectric interaction? | Is the majority of biological damage from X-rays a direct result of the initial photon-induced ionization? | Does the probability of the photoelectric effect increase with increasing incident photon energy? | Is the kinetic energy of the ejected photoelectron the sum of the incident photon's energy and the electron's binding energy? | FALSE | For the effect to occur, the photon energy must be greater than the electron's binding energy. The probability is highest when the photon energy is just above the binding energy (e.g., the K-edge). | TRUE | The ejection of the photoelectron leaves a vacancy in an inner shell. An outer-shell electron fills this vacancy, releasing energy as either characteristic radiation or an Auger electron. | FALSE | The majority of biological damage is caused indirectly. The high-energy photoelectron created goes on to produce thousands of additional ionizations along its track. | FALSE | The probability of the photoelectric effect is inversely proportional to the cube of the energy (1/E³). As photon energy increases, the likelihood of this interaction decreases rapidly. | FALSE | The kinetic energy (KE) of the photoelectron is the incident photon's energy (E) minus the electron's binding energy (BE): KE = E - BE.
Pertaining to the interactions between X-rays and matter: | Is the photoelectric effect the dominant interaction in high-energy techniques like computed tomography (CT)? | For photon energies of 30 keV and below, is Compton scattering the most probable interaction in tissue? | In high-kV diagnostic imaging, such as CT, is pair production a significant interaction mechanism? | Is the recoil electron resulting from a Compton scatter event ever scattered in a backward trajectory? | Is the likelihood of a Compton scatter interaction dependent on the electron density of the absorbing material? | FALSE | In soft tissue, Compton scatter is the dominant interaction for the energies used in CT (e.g., effective energies of 60-80 keV). The photoelectric effect dominates at lower energies. | FALSE | In soft tissue at ~26 keV, the probabilities are roughly equal. Below this energy, the photoelectric effect becomes the dominant interaction. | FALSE | Pair production requires a minimum photon energy of 1.02 MeV, which is far above the diagnostic imaging range. It is relevant only in radiation therapy. | FALSE | Conservation of energy and momentum dictates that the Compton recoil electron is always scattered in a forward direction, at an angle between 0° and 90° to the incident photon's path. | TRUE | The probability of a Compton interaction is directly proportional to the number of electrons per unit volume (electron density), which is closely related to the physical density of the material.
On the topic of luminescence phenomena: | In a luminescent material, is the highest-energy electron band that is normally filled referred to as the conduction band? | Does the phenomenon of luminescence require a material lattice of perfect purity, devoid of any activator impurities? | For phosphors used in radiographic detectors, is the intensity of emitted light directly proportional to the amount of absorbed X-ray or gamma-ray energy? | Does the image readout process in computed radiography (CR) involve heating the exposed phosphor plate to stimulate the release of light? | Do optically stimulated luminescent (OSL) dosimeters offer superior measurement sensitivity compared to thermoluminescent detectors (TLDs)? | FALSE | The highest normally filled electron band is the valence band. The conduction band is the normally empty band above it, to which electrons are excited. | FALSE | Practical phosphors rely on small, controlled amounts of impurities, known as "activators," which create the electron traps within the band gap that are essential for luminescence. | TRUE | In an ideal phosphor, the amount of light produced (scintillation) is directly proportional to the energy deposited by the incident radiation, allowing for accurate energy measurement. | FALSE | CR readout uses photostimulable luminescence, where a red laser scans the plate to stimulate the release of stored energy as blue-green light. Heating is used for thermoluminescent dosimeters (TLDs). | TRUE | OSL dosimeters are more sensitive, allowing for measurement of lower radiation doses than TLDs. They also have the key advantage of being able to be re-read multiple times.
General principles of radiation: | Does the energy spectrum of Bremsstrahlung radiation extend continuously from near zero up to the peak voltage potential in the X-ray tube? | Does K-alpha (Kα) radiation refer to the emission of an alpha particle? | Does the process of characteristic radiation yield photons with discrete energies that correspond to the electron binding energies of the target atom? | Is L-shell characteristic radiation from a tungsten target diagnostically significant in high-kV imaging techniques? | Is beta particle emission a frequently utilized form of radiation in nuclear medicine diagnostic procedures? | TRUE | Bremsstrahlung radiation forms a continuous spectrum of energies, ranging from near zero up to a maximum value determined by the peak kilovoltage (kVp) applied across the tube. | FALSE | K-alpha radiation is a type of characteristic X-ray produced when an L-shell electron fills a K-shell vacancy. An alpha particle is a helium nucleus (2 protons, 2 neutrons). | TRUE | The energy of a characteristic photon is equal to the difference in binding energies between the two electron shells involved, making these energies unique, or "characteristic," of the target element. | FALSE | L-shell X-rays from tungsten have energies of about 8-12 keV. They are too low-energy to pass through the X-ray tube's inherent filtration and do not contribute to the final image. | FALSE | Beta particles (electrons) have a very short range in tissue and cannot be imaged externally. Gamma rays are used for diagnostic imaging (SPECT/PET), while beta emitters are mainly used for therapy.
Which of the following statements regarding protons are correct? | They have a negative charge | They are equal to the number of electrons in a non-ionized atom | They are equal to the atomic number in a non-ionized atom | They have no mass | They increase in number relative to neutrons, following electron capture | FALSE | Protons have a positive charge. Electrons have a negative charge. | TRUE | In a neutral (non-ionized) atom, the number of positive charges (protons) must equal the number of negative charges (electrons). | TRUE | The atomic number (Z) of an element is defined by its number of protons. | FALSE | Both protons and neutrons have mass and contribute to the atomic mass number of an atom. Electrons have a negligible mass in comparison. | FALSE | Following electron capture, a proton combines with an electron to form a neutron (p + e -> n), thus decreasing the number of protons relative to neutrons.
The binding energy of electrons: | Is the energy expended in moving an electron from an inner to an outer shell | Is higher for an L-shell electron than an M-shell electron | Is influenced by the number of neutrons within an atom | Determines Bremsstrahlung photon energy | Determines the energy of the photoelectron produced following photoelectric absorption | FALSE | Binding energy is the energy required to completely remove an electron from the atom against the electrostatic pull of the nucleus. | TRUE | Inner shells (like the L-shell) are closer to the positive nucleus and are more tightly bound than outer shells (like the M-shell), thus having a higher binding energy. | FALSE | Neutrons are electrically neutral and do not influence the electrostatic force that determines the binding energy of electrons. | FALSE | Bremsstrahlung energy is determined by the deceleration of incident electrons near the target nucleus. Characteristic radiation energy is determined by electron binding energies. | TRUE | The kinetic energy of the ejected photoelectron is equal to the incident photon's energy minus the binding energy of the electron.
Which of the following are correct regarding electromagnetic radiation? | Gamma rays are a form of electromagnetic radiation | The particles have a mass equivalent to that of neutrons | In a vacuum, the velocity of the particles differs depending on their individual properties | It results in a sinusoidal graph when magnetic field strength is plotted against time | The frequency is the interval between two successive crests | TRUE | Electromagnetic radiation includes radio waves, X-rays, and gamma rays, distinguished by their origin and energy. | FALSE | Electromagnetic radiation consists of photons, which have no mass or electric charge. | FALSE | All forms of electromagnetic radiation travel at the same speed in a vacuum: the speed of light (c). | TRUE | An electromagnetic wave consists of oscillating electric and magnetic fields that are perpendicular to each other, which can be represented by a sinusoidal graph. | FALSE | Frequency (f) is the number of wave crests passing a point per second. The interval between successive crests is the period (T).
Regarding electromagnetic beam radiation: | Energy fluence is the number of photons per unit area of a beam | Beam intensity is the total amount of energy per unit area travelling per unit time | Wavelength is inversely proportional to frequency | Frequency is inversely proportional to photon energy (keV) | Photon energy is inversely proportional to wavelength | FALSE | The number of photons per unit area is photon fluence. The sum of the energies of those photons per unit area is energy fluence. | TRUE | This is the correct definition of beam intensity (or energy flux density). | TRUE | This is a fundamental wave relationship, described by the equation c = λf, where c is the constant speed of light. | FALSE | Photon energy (E) is directly proportional to frequency (f), as described by Planck's equation, E = hf, where h is Planck's constant. | TRUE | Since energy is proportional to frequency, and frequency is inversely proportional to wavelength, it follows that energy is inversely proportional to wavelength.
Concerning electromagnetic rays: | They travel parallel to each other in a straight line | Beam intensity is proportional to the square of the amplitude | The area of the beam is proportional to the square of the distance as it travels away from a point source | Beam intensity is proportional to the square of the distance as it travels away from a point source | The inverse square law is useful for estimating the intensity of a beam once it has passed through a copper filter | FALSE | Rays from a point source diverge as they travel away in straight lines; they are not parallel. | TRUE | The intensity of a wave is proportional to the square of its amplitude. | TRUE | As the beam diverges from a point source, the area it covers increases with the square of the distance. | FALSE | Beam intensity is inversely proportional to the square of the distance, according to the inverse square law. | FALSE | The inverse square law only applies in non-attenuating conditions. A filter, by definition, attenuates the beam, making the law inapplicable.
Which of the following statements are correct regarding an X-ray tube? | The tube current (mA) is increased by increasing the filament voltage | An increase in the tube voltage (kV) leads to a proportional increase in tube current (mA) | Electrons are generated by heating the anode, which is usually made from tungsten | The kinetic energy of electrons (keV) in the X-ray tube is dependent on the tube voltage (kV) | The collision of electrons with a tungsten target mainly results in the production of X-ray radiation | TRUE | Increasing the voltage/current to the filament increases its temperature, which increases the rate of thermionic emission, and thus the number of electrons (tube current, mA). | FALSE | The tube voltage (kV) determines the kinetic energy of each electron, not the number of electrons (tube current, mA). | FALSE | Electrons are generated by heating the filament at the cathode. | TRUE | The tube voltage (kV) creates the potential difference that accelerates the electrons, so their maximum kinetic energy (in keV) is numerically equal to the kV. | FALSE | X-ray production is highly inefficient. Over 99% of the kinetic energy of the electrons is converted into heat upon collision with the target.
Which of the following are true regarding the production of X-ray photons when using a tungsten target? | The majority of X-rays emitted are a result of characteristic radiation | The kinetic energy of electrons (keV) interacting with the target is equal to the kV between the anode and cathode of the X-ray tube | Filament electrons with 90 keV can dislodge K-shell electrons in the target | The energy of Ka radiation is greater than Kb radiation | L-shell radiation makes up 25% of characteristic radiation emitted from the tube | FALSE | Approximately 80% of the X-rays produced in a typical diagnostic X-ray tube are from Bremsstrahlung radiation. | TRUE | The peak kinetic energy of electrons in keV is numerically equal to the peak tube voltage in kV. | TRUE | To dislodge a K-shell electron from tungsten, the incident electron must have energy greater than tungsten's K-shell binding energy, which is ~70 keV. 90 keV is sufficient. | FALSE | Kα radiation (L to K shell transition) has less energy than Kβ radiation (M to K shell transition) because the energy gap between L and K is smaller than between M and K. | FALSE | Characteristic X-rays from L-shell interactions in tungsten have very low energy (~12 keV) and are absorbed by the X-ray tube's own filtration, so they do not exit the tube.
Concerning characteristic radiation: | It results in photons with a fixed energy, for a given material | It mainly involves filament electrons dislodging L-shell electrons | Photon energy is directly proportional to the tube voltage | The rate of production of characteristic radiation is directly proportional to the filament voltage | Atomic number influences photon energy of K-radiation | TRUE | The photon energy is equal to the difference between two specific electron shell binding energies, which is constant and "characteristic" for a given target element. | FALSE | For characteristic X-rays to be useful in diagnostic imaging (e.g., from tungsten), the interaction must involve dislodging high-binding-energy K-shell electrons. | FALSE | The tube voltage determines if characteristic radiation can be produced (must be > binding energy), but it does not influence the fixed energy of the resulting photons. | TRUE | Increasing the filament voltage/current increases the tube current (mA), which increases the number of electrons hitting the target, thus proportionally increasing the production rate of all X-rays. | TRUE | The binding energy of electron shells increases with the atomic number (Z). Therefore, the energy of K-radiation is dependent on the target material's Z.
Concerning Bremsstrahlung radiation: | It results mainly from filament electrons colliding with the nucleus | It can result in photons that are equivalent in energy to the tube voltage | It results in X-ray photons of the same energy, which are dependent on the target material | An increase in atomic number of the target material results in an increase in the energy of photons | A reduction in atomic number results in reduced photon production | FALSE | It results from filament electrons being deflected and slowed down by the electrostatic field of the nucleus. A direct collision is extremely rare. | TRUE | In the rare event that an electron is completely stopped by its interaction with the nucleus, it loses all its kinetic energy, which is emitted as a single photon with energy equal to the kV. | FALSE | Bremsstrahlung radiation produces a continuous spectrum of photons with varying energies, from near zero up to the maximum kinetic energy of the incident electron. | FALSE | Increasing the atomic number (Z) of the target increases the efficiency of production (number of photons), but does not affect the maximum possible photon energy, which is set by the kV. | TRUE | The efficiency of Bremsstrahlung production is proportional to the atomic number (Z) of the target. A lower Z target will produce fewer photons.
Which of the following are true with regard to X-ray photon production? | The area of a continuous spectrum graph represents the total output of all X-ray photons emitted as a result of characteristic radiation | Increasing the tube voltage (kV) results in the continuous spectrum shifting to the right and the line spectrum increasing in height | One-quarter of the photons from Bremsstrahlung radiation that reach the patient are less than 20 keV | Increasing the filament voltage causes Bremsstrahlung and characteristic radiation graphs to increase in height | Reducing the filament voltage has no effect on the maximum photon energy produced by characteristic or Bremsstrahlung radiation | FALSE | The area under the continuous spectrum represents the total output from Bremsstrahlung radiation. Characteristic radiation appears as discrete peaks (line spectrum). | TRUE | Increasing kV increases the maximum energy (shifts spectrum right) and makes production more efficient (increases height of continuous spectrum and characteristic peaks). | FALSE | Photons with energy less than ~20 keV are generally absorbed by the inherent and added filtration of the X-ray tube and do not reach the patient. | TRUE | Increasing filament voltage increases tube current (mA), which increases the number of electrons. This raises the quantity of both types of radiation, increasing the height of the entire spectrum. | TRUE | Filament voltage/current controls the quantity (number) of electrons, not their energy. The maximum photon energy is determined solely by the tube voltage (kV).
Concerning the interaction of X-ray photons with matter: | It is possible to predict the fraction of photons that will be absorbed or scattered, when passing through a given material | Attenuation is represented by the number of photons absorbed or scattered by matter | Scattered photons help form the primary image on the film | The half-value layer (HVL) of a material is the thickness needed to reduce the number of photons in a beam by half | The HVL is inversely proportional to the linear attenuation coefficient (LAC) of a material | FALSE | Individual photon interactions are random (stochastic). While we can predict the behavior of a large number of photons, we cannot predict the fate of any single one. | TRUE | Attenuation is the total reduction in the number of primary photons in the beam due to both absorption and scattering interactions. | FALSE | Transmitted (unscattered) photons form the useful image. Scattered photons add unwanted noise (fog) that degrades image contrast. | FALSE | HVL is the thickness required to reduce the intensity (energy fluence rate) of a monoenergetic beam by half, not necessarily the number of photons. | TRUE | The LAC (µ) represents the probability of interaction per unit length. The relationship is HVL = 0.693 / µ. A high LAC means a low HVL.
Concerning factors affecting absorption and scatter of X-rays: | An increase in the density of a material results in an increase in the linear attenuation coefficient (LAC) | Increasing the atomic number of a material increases its half-value layer (HVL) | Reducing the tube voltage increases the LAC | Increasing the filament voltage reduces the HVL of a material | A narrow beam has a higher scatter-to-transmission ratio than a wide beam | TRUE | A denser material has more atoms/electrons per unit volume, increasing the probability of interaction for a photon passing through it, thus increasing the LAC. | FALSE | A higher atomic number greatly increases the probability of photoelectric absorption, leading to more attenuation. This reduces the HVL. | TRUE | Reducing the tube voltage lowers the average photon energy. Lower energy photons are more likely to be attenuated (especially via photoelectric effect), thus the LAC increases. | FALSE | Filament voltage affects the quantity of photons, not their energy or penetrating power. Therefore, it has no effect on the HVL. | FALSE | A narrow beam irradiates a smaller volume of tissue, which generates less scatter. Therefore, a narrow beam has a lower scatter-to-transmission ratio.
Concerning an X-ray beam: | It reduces in equal quantities as it passes through material of equal thickness | No matter how thick the material, it is not possible to completely absorb the primary X-ray beam | Beam hardening results from a reduced number of photons being removed from the primary beam | For a heterogeneous beam, the half-value layer (HVL) increases as the beam passes through the material | Wearing lead gloves protects the operator’s fingers from the X-ray beam | FALSE | This is only true for a mono-energetic beam. A clinical (poly-energetic) beam gets "harder" as it passes through material, so subsequent layers attenuate it less. | TRUE | Attenuation is an exponential process. The intensity approaches zero asymptotically but never mathematically reaches it. | FALSE | Beam hardening results from low-energy photons being preferentially removed from the beam, which increases the average energy of the remaining photons. | TRUE | As the beam hardens (average energy increases), its penetrating power increases. A greater thickness of material is then required to halve its intensity, so the HVL increases. | FALSE | Protective equipment like lead gloves and aprons protects from scattered radiation, not the much more intense primary beam.
Which of the following are correct with regard to attenuation of X-ray photons? | The Compton effect refers to the interaction of X-ray photons with free electrons | Attenuation is the difference between the incident beam and the attenuated beam | Photoelectric absorption refers to the interaction of X-ray photons with loosely bound electrons | An increase in atomic number increases the linear attenuation coefficient (LAC) | Elastic scatter results in no loss of energy | TRUE | The Compton effect involves an incident photon interacting with an outer-shell electron that is so loosely bound it can be considered free. | TRUE | The attenuated beam is what emerges after the incident beam has passed through matter. The difference represents the photons that were attenuated (absorbed or scattered). | FALSE | Photoelectric absorption involves the interaction of a photon with a tightly bound, inner-shell electron (e.g., K-shell). | TRUE | An increase in atomic number (Z) dramatically increases the probability of photoelectric effect (proportional to Z³), which significantly increases the overall LAC. | TRUE | Elastic (also known as coherent or Rayleigh) scatter involves the photon changing direction without transferring any energy to the atom; hence, there is no loss of energy.
Concerning Compton interaction, which of the following are correct? | It involves the collision of X-ray photons with any electron | The probability that Compton attenuation will occur decreases as photon energy is increased | Increasing the tube voltage results in a higher proportion of side scatter | The higher the scatter angle, the greater the penetration of the recoil electrons | An increase in the incident photon energy results in scatter photons with greater energy | FALSE | Compton interaction specifically involves photons interacting with free or loosely bound outer-shell electrons. | TRUE | While the decrease is not as rapid as with the photoelectric effect, the probability of Compton interaction does generally decrease as photon energy increases through the diagnostic range. | FALSE | Higher energy incident photons result in more forward-directed scatter. Therefore, increasing the tube voltage (and thus photon energy) leads to less side and backscatter. | TRUE | As the photon's scatter angle increases, it transfers more energy to the recoil electron. A more energetic electron has greater penetration. | TRUE | While the scattered photon always has less energy than the incident photon, a higher energy incident photon will generally result in a higher energy scattered photon (for a given angle).
Regarding fundamental radiation quantities and units: | The absorbed dose is a measure of biological risk and is expressed in Sieverts (Sv) | Equivalent dose accounts for the different sensitivities of various body tissues to radiation | The SI unit for absorbed dose is the Gray (Gy) | For diagnostic X-rays, the effective dose is numerically identical to the equivalent dose | Exposure is a measure of the ionization produced by photons in human tissue | FALSE | Absorbed dose measures the energy deposited per unit mass and is expressed in Gray (Gy). The Sievert (Sv) is used for equivalent and effective dose. | FALSE | Equivalent dose accounts for the different biological effectiveness of various types of radiation (using the radiation weighting factor, wR​). Effective dose accounts for tissue sensitivity. | TRUE | The Gray (Gy) is the SI unit for absorbed dose, defined as one joule of energy absorbed per kilogram of matter (1 J/kg). | FALSE | While the radiation weighting factor (wR​) for X-rays is 1, the effective dose is calculated by summing the equivalent doses to all tissues, each multiplied by its tissue weighting factor (wT​). | FALSE | Exposure is a quantity that specifically measures the ionization produced by X-rays or gamma rays in a unit mass of air.
Concerning the ALARA principle: | The acronym ALARA stands for "All Levels Are Readily Achievable" | It is a regulatory requirement that mandates radiation doses be kept below a specific, fixed threshold | It applies to occupational and public exposure but not to patient exposure | The principle implies that no dose of radiation can be considered completely safe | The "Reasonably" component suggests that social and economic factors may be considered when implementing protection measures | FALSE | ALARA stands for "As Low As Reasonably Achievable." It is the cornerstone of modern radiation protection philosophy. | FALSE | ALARA is a guiding principle to minimize dose, not a fixed threshold. Dose limits are the legal maximums, but practice should always be guided by ALARA. | FALSE | The ALARA principle applies to all radiation exposures, including those to patients, staff, and the public. For patients, it means using the minimum dose required to achieve a diagnostic result. | TRUE | Because of the linear, no-threshold model for stochastic effects like cancer, ALARA operates on the assumption that even the smallest dose carries some risk, which must be managed. | TRUE | This is a key aspect. Radiation protection measures should be optimized to provide the greatest benefit, considering the costs and practicality of implementation.
Regarding the cardinal principles of radiation protection: | The three principles are Time, Distance, and Collimation | Doubling the exposure time will reduce the total dose received by half | Shielding protects individuals by absorbing all incident radiation photons completely | Distance is often the most effective method for reducing external radiation exposure | The effectiveness of shielding is independent of the energy (kVp) of the X-ray beam | FALSE | The three cardinal principles are Time, Distance, and Shielding. Collimation is an important dose reduction technique but not one of the three core principles. | FALSE | Radiation dose is directly proportional to exposure time. Doubling the time will double the dose received. | FALSE | Shielding works by attenuation, which reduces the intensity of radiation by absorbing or scattering a fraction of the photons. It does not stop all radiation. | TRUE | Due to the inverse square law, doubling the distance from a point source reduces the dose rate to one-quarter, making it a very powerful protection tool. | FALSE | Higher energy photons are more penetrating and require thicker or denser shielding to achieve the same level of protection. The shielding's effectiveness is highly energy-dependent.
Concerning deterministic and stochastic effects of radiation: | Cancer induction is considered a deterministic effect | Deterministic effects have a dose threshold below which they do not occur | The severity of a stochastic effect is proportional to the radiation dose received | Skin erythema (reddening) is a common stochastic effect of radiation | Stochastic effects are characterized by a delayed onset, often appearing years after exposure | FALSE | Cancer induction is the primary example of a stochastic effect, where the probability of occurrence, not the severity, increases with dose. | TRUE | Deterministic (or tissue) effects, such as cataracts or skin burns, only appear after a certain threshold dose has been exceeded. | FALSE | For a stochastic effect, the probability of the effect occurring increases with dose. The severity of the effect (e.g., the aggressiveness of a cancer) is independent of the dose. | FALSE | Skin erythema is a classic deterministic effect that occurs only after a high threshold dose (typically >2 Gy) has been delivered to the skin. | TRUE | This is a hallmark of stochastic effects like radiation-induced cancers, which may have a latency period of many years or even decades.
Regarding personnel dosimetry and dose limits: | The annual effective dose limit for an occupational radiation worker is 5 mSv | A declared pregnant worker has a monthly dose limit to the fetus of 5 mSv | Personnel dosimeters, like TLDs or OSLs, provide immediate real-time dose readings | The dosimeter should be worn on the trunk of the body, underneath a lead apron | The annual equivalent dose limit for the lens of the eye for a worker is the same as the whole-body limit | FALSE | The annual effective dose limit for an occupational worker is 20 mSv, averaged over 5 years, with no single year exceeding 50 mSv. | FALSE | The dose limit to the surface of the abdomen of a declared pregnant worker is 2 mSv for the remainder of the pregnancy, and the fetal dose limit is 1 mSv. | FALSE | TLDs and OSLs are passive dosimeters that accumulate dose over a period (e.g., a month or quarter) and must be sent away for processing. Electronic dosimeters provide real-time readings. | FALSE | When a lead apron is worn, the primary dosimeter should be worn on the collar, outside the apron, to measure the dose to the more sensitive, unshielded head and lens of the eye. | FALSE | The annual equivalent dose limit for the lens of the eye for an occupational worker is 20 mSv, which is different from the effective dose limit for the whole body.
Concerning shielding in a diagnostic radiology department: | The primary protective barrier is designed to attenuate the useful (primary) X-ray beam | Lead (Pb) is used for shielding primarily because it is the densest element available | The half-value layer (HVL) of lead is constant for all X-ray beam energies | A lead apron with 0.5 mm lead equivalence will attenuate approximately 75% of a 100 kVp X-ray beam | Secondary barriers protect against leakage and primary radiation | FALSE | The primary barrier is designed to attenuate the useful X-ray beam. | TRUE | Lead has a high atomic number (Z=82) and high density, which makes it extremely effective at attenuating X-rays via the photoelectric effect. | FALSE | The HVL is highly dependent on photon energy. A higher energy (higher kVp) beam is more penetrating and will have a greater HVL in lead than a lower energy beam. | FALSE | A 0.5 mm lead apron will attenuate approximately 90% or more of a typical diagnostic beam. A 0.25 mm apron attenuates ~75% at 100 kVp. | FALSE | Secondary barriers are designed to protect against scatter and leakage radiation only, which are much less intense than the primary beam.
Regarding patient dose reduction techniques: | Using a low-kVp and high-mAs technique generally reduces the patient's absorbed dose | Increasing the source-to-skin distance (SSD) increases the patient entrance dose | Beam filtration is used to remove high-energy photons from the X-ray beam | Collimating the beam to the smallest necessary area is a key method for reducing dose and scatter | Grids are used to reduce patient dose by absorbing scattered radiation | FALSE | A high-kVp and low-mAs technique is preferred, as it produces a more penetrating beam, which reduces the patient's absorbed dose while maintaining image quality. | FALSE | Increasing the SSD decreases the patient's entrance skin dose due to the inverse square law. | FALSE | Filtration (e.g., aluminum) is used to absorb low-energy ("soft") X-rays that would contribute to patient skin dose without contributing to the image. | TRUE | Restricting the beam to the area of interest reduces the total volume of tissue irradiated, which is the single most effective way to reduce overall patient dose and scattered radiation. | FALSE | While grids improve image quality by absorbing scatter, they necessitate an increase in exposure factors (mAs), which increases the patient's radiation dose.
Pertaining to radiation weighting factors (wR​): | The radiation weighting factor for alpha particles is 1 | A 1 Gy dose of alpha particles results in a lower equivalent dose than a 1 Gy dose of X-rays | wR​ is used to convert effective dose into equivalent dose | The purpose of wR​ is to account for the fact that different types of radiation cause different amounts of biological damage for the same absorbed dose | Beta particles have a higher wR​ than gamma rays | FALSE | Alpha particles are densely ionizing and cause significant biological damage. Their radiation weighting factor (wR​) is 20. | FALSE | Because the wR​ of alpha particles is 20, a 1 Gy dose results in an equivalent dose of 20 Sv, which is much higher than the 1 Sv resulting from a 1 Gy dose of X-rays (wR​=1). | FALSE | wR​ is used to convert absorbed dose (Gy) into equivalent dose (Sv). Equivalent Dose = Absorbed Dose x wR​. | TRUE | This is the fundamental purpose of the radiation weighting factor. It normalizes risk across different radiation types. | FALSE | Both beta particles (electrons) and gamma rays (photons) are sparsely ionizing radiations and are assigned a radiation weighting factor (wR​) of 1.
Regarding tissue weighting factors (wT​): | The tissue weighting factor reflects the physical density of an organ | Tissues that are highly radiosensitive, like the gonads, have a low tissue weighting factor | The sum of all tissue weighting factors for the entire body is equal to 100 | wT​ is used to calculate the effective dose from the absorbed dose | The skin has a higher tissue weighting factor than the bone marrow | FALSE | The tissue weighting factor (wT​) reflects the relative radiosensitivity of an organ or tissue in contributing to the overall stochastic risk (cancer and hereditary effects). | FALSE | Highly radiosensitive tissues (those that contribute most to overall risk), such as the gonads, red bone marrow, and colon, have the highest tissue weighting factors. | FALSE | The sum of all tissue weighting factors is, by definition, equal to 1.0. | FALSE | wT​ is used to calculate the effective dose from the equivalent doses to various tissues. Effective Dose = Σ (wT​ × Equivalent Dose). | FALSE | The skin has one of the lowest tissue weighting factors (wT​=0.01), while the red bone marrow has one of the highest (wT​=0.12), reflecting its high sensitivity.
Concerning scatter radiation in diagnostic imaging: | Scatter radiation originates primarily from coherent scattering interactions | The amount of scatter produced is inversely proportional to the patient thickness | Scatter radiation is the primary source of occupational radiation exposure for radiologic technologists | Scatter radiation travels predominantly in the forward direction, parallel to the primary beam | Using an air gap technique reduces the amount of scatter produced in the patient | FALSE | Scatter radiation in the diagnostic energy range originates almost entirely from Compton scattering interactions. | FALSE | The amount of scatter produced is directly proportional to the volume of tissue irradiated, so thicker patients produce more scatter. | TRUE | Technologists are positioned outside the primary beam, so the vast majority of the dose they receive comes from X-rays scattered from the patient. | FALSE | Compton-scattered photons can be deflected in any direction. Side and backscatter are major contributors to staff dose. | FALSE | The air gap technique does not reduce scatter production; it works by allowing scatter, which travels at an angle, to miss the detector, thereby improving image contrast.
Regarding pregnancy and radiation exposure: | The period of greatest risk for the induction of congenital malformations is the first two weeks of gestation | A fetal dose of 500 mGy is likely to result in severe intellectual disability | The dose limit for a declared pregnant worker ensures the fetal dose will not exceed 10 mSv | A diagnostic X-ray of the abdomen is absolutely contraindicated in any pregnant patient | The primary radiation risk to the fetus during the first trimester is the induction of childhood cancer | FALSE | The first two weeks is the "all-or-nothing" period, where a significant dose may cause embryonic death but not malformations. The period of greatest risk for malformations is weeks 2 through 8. | TRUE | A high dose (> 0.5 Gy or 500 mGy) during the period of major organogenesis and brain development (weeks 8-25) has a high probability of causing severe intellectual disability and other congenital abnormalities. | FALSE | The dose limit to the abdomen of the worker is set to ensure the total equivalent dose to the fetus over the course of the pregnancy does not exceed 1 mSv. | FALSE | No diagnostic X-ray is absolutely contraindicated. The decision is based on a risk-benefit analysis, and if the exam is clinically justified, it should be performed using dose-reduction techniques. | FALSE | While there is a small risk of cancer induction, the primary concern during the first trimester (specifically weeks 2-8) is teratogenesis (the production of malformations).
Pertaining to leakage radiation and controlled areas: | Leakage radiation is another term for scattered radiation from the patient | The regulatory limit for leakage radiation from an X-ray tube housing is 1 mGy per hour at a distance of 1 meter | A "Controlled Area" is an area accessible to the general public | The shielding design for a controlled area aims to keep the dose to occupants below 1 mSv per week | Occupational workers receive most of their dose from leakage radiation | FALSE | Leakage radiation is radiation that escapes through the protective X-ray tube housing, not scatter from the patient. | TRUE | This is the standard regulatory limit for diagnostic X-ray tube housings when operated at their maximum continuous ratings. | FALSE | A "Controlled Area" is one that is restricted and supervised for the purpose of radiation protection. An area for the general public is an "Uncontrolled Area." | FALSE | Shielding for a controlled area is designed to keep the dose to a radiation worker below the occupational limit, typically aiming for less than 0.1 mSv per week. | FALSE | Occupational workers receive the vast majority of their dose from scatter radiation originating from the patient. Leakage radiation is typically a very minor component.
Regarding radiation effects on the lens of the eye: | Radiation-induced cataracts are considered a stochastic effect | The lens of the eye is considered highly resistant to radiation damage | The dose threshold for detectable lens opacities is approximately 0.5 Gy (500 mGy) | The annual equivalent dose limit for the lens for the public is 50 mSv | Wearing leaded eyewear can significantly reduce the dose to the lens during fluoroscopy | FALSE | Radiation-induced cataracts are a classic example of a deterministic (tissue) effect, as a threshold dose must be exceeded for them to occur. | FALSE | The lens of the eye is now known to be one of the more radiosensitive tissues in the body, which is why its specific dose limit was recently lowered by regulatory bodies. | TRUE | Recent evidence has established that the threshold for radiation-induced cataracts is about 0.5 Gy, which is much lower than previously thought. | FALSE | The annual equivalent dose limit for the lens of the eye for a member of the general public is 15 mSv. | TRUE | Leaded eyewear, particularly wrap-around styles, can reduce the dose to the lens from scatter radiation by 90% or more and is strongly recommended for interventional staff.
Concerning quality control and radiation safety: | A benefit of a Quality Assurance (QA) program is the reduction of patient and staff doses | Linearity checks ensure that the mAs setting is consistent regardless of the mA and time selected | The HVL of the X-ray beam is measured to ensure adequate filtration | Reproducibility means that the X-ray output should be the same for consecutive exposures at the same technique factors | Collimator accuracy tests ensure that the lead shutters are functioning correctly | TRUE | A key goal of QA is to ensure optimal image quality at the minimum necessary dose, which benefits both patients and staff by reducing repeat exposures and unnecessary radiation. | FALSE | This describes reciprocity. Linearity checks ensure that the radiation output (mGy/mAs) remains constant when the kVp is fixed but the mA and time are changed to produce the same mAs. | TRUE | Measuring the half-value layer (HVL) is the most practical method to assess the overall quality (penetrability) of the X-ray beam and confirm that the total filtration is adequate. | TRUE | A key QC test for the generator, reproducibility requires that the radiation output be consistent, typically within a ±5% variance, for repeated exposures at identical settings. | TRUE | These tests verify that the light field shown by the collimator accurately corresponds to the actual X-ray field, typically requiring alignment within 2% of the SID.
Regarding various sources of radiation exposure: | The largest component of natural background radiation exposure for the average person is from cosmic rays | A typical chest X-ray delivers an effective dose of approximately 5 mSv | Medical imaging is now the largest source of man-made radiation exposure to the U.S. population | A coast-to-coast airplane flight results in a negligible radiation dose | Living in a brick or stone house can increase one's background radiation exposure | FALSE | The largest single contributor to natural background radiation is from inhaling naturally occurring radon gas and its decay products, not cosmic rays. | FALSE | A typical posterior-anterior (PA) chest X-ray delivers a very low effective dose, on the order of 0.02 mSv. 5 mSv is equivalent to a CT scan of the abdomen. | TRUE | The use of medical imaging, particularly computed tomography (CT), has grown significantly and now surpasses all other man-made sources of exposure for the general population. | FALSE | A coast-to-coast flight results in a dose of approximately 0.03-0.05 mSv due to increased exposure to cosmic radiation at high altitudes. This is more than a chest X-ray. | TRUE | Building materials like brick, stone, and concrete contain naturally occurring radioactive elements (e.g., uranium, thorium), which can slightly increase the background radiation dose compared to living in a wooden house.
Regarding the photoelectric effect: | It results in the production of Bremsstrahlung radiation | It results in the production of scattered photons, the energy of which is dependent on the initial photon energy (keV) | X-rays passing through barium cause greater scatter than those passing through human tissues | Ejection of a K-shell electron by an incident photon results in the production of an Auger electron | Auger electrons are produced as an indirect result of photoelectric radiation | FALSE | The photoelectric effect leads to the production of characteristic radiation (or Auger electrons) when the resulting electron vacancy is filled. Bremsstrahlung is from electron deceleration. | FALSE | The photoelectric effect involves the complete absorption of the incident photon; no scattered photon is produced. Compton scatter produces scattered photons. | FALSE | Barium has a high atomic number, which causes significantly greater photoelectric absorption, not greater scatter. Soft tissue interactions are dominated by Compton scatter. | FALSE | It results in a vacancy. The filling of this vacancy can produce either a characteristic X-ray or an Auger electron. It is not the only possible outcome. | TRUE | The initial photoelectric interaction creates an inner-shell vacancy. The subsequent de-excitation of the atom can result in the emission of an Auger electron.
Concerning factors influencing attenuation: | In material with a low atomic number, the photoelectric effect results in complete absorption | The energy of the incident photon must be greater than the binding energy of the electron in photoelectric absorption | An increase in the incident photon energy results in recoil electrons with lower penetration | Reducing the tube voltage (kV) results in less scatter reaching the film | The kinetic energy of the photoelectron ejected as a result of the photoelectric effect is not dependent on the atomic number of the material | TRUE | The photoelectric effect, by definition, is a complete absorption event, regardless of the material's atomic number (Z). The probability of the effect is just lower in low-Z materials. | TRUE | This is a fundamental requirement for the photoelectric effect to occur. The photon must have enough energy to overcome the electron's binding energy and eject it from its shell. | FALSE | This refers to Compton scatter. A higher energy incident photon transfers more energy to the recoil electron, giving it higher kinetic energy and thus greater penetration. | TRUE | Lowering the kVp reduces the average photon energy. This increases the probability of the photoelectric effect relative to Compton scatter, thus generating less scattered radiation overall. | FALSE | The photoelectron's kinetic energy is equal to the incident photon energy minus the electron's binding energy. Since binding energy is highly dependent on atomic number (Z), so is the photoelectron's energy.
Which of the following are true regarding the linear attenuation coefficient (LAC)? | The photoelectric LAC is directly proportional to photon energy | The total LAC is dependent on the physical density of the material | The Compton LAC is dependent on atomic number | The photoelectric LAC increases as the atomic number increases | The Compton LAC is dependent on the electron density | FALSE | The photoelectric LAC is inversely proportional to approximately the cube of the photon energy (1/E³). | TRUE | The LAC is directly proportional to the physical density of the material, as denser materials contain more atoms per unit volume for photons to interact with. | FALSE | The Compton effect is nearly independent of the atomic number (Z). Its probability depends on the electron density of the material. | TRUE | The photoelectric LAC is strongly dependent on the atomic number, approximately proportional to its cube (Z³). | TRUE | The probability of a Compton interaction is directly proportional to the number of electrons per unit volume, which is known as the electron density.
Concerning absorption edges: | K-edge binding energy is lower than L-edge binding energy | Between the K-shell and L-shell, the increase in photoelectric attenuation is proportional to the photon energy | For tungsten, the K-shell binding energy (Ek) is equal to 74 keV | There is a sudden increase in attenuation when photon energy reaches L-shell binding energy (EL) | When choosing a filter, it is important to make sure that the peak of the X-ray spectrum lies on the high-energy side of its absorption edge | FALSE | The K-shell is the innermost electron shell and has the highest binding energy. Binding energy decreases for outer shells (K > L > M). | FALSE | Between absorption edges, the photoelectric attenuation decreases rapidly as photon energy increases (it is proportional to 1/E³). | FALSE | The atomic number (Z) of tungsten is 74. Its K-shell binding energy is approximately 69.5 keV. | TRUE | At any absorption edge (K, L, M, etc.), when the photon energy becomes just sufficient to eject an electron from that shell, a new interaction channel opens, causing a sharp increase in attenuation. | TRUE | This describes a K-edge filter. The filter strongly absorbs photons with energies just above its K-edge, effectively creating a more monoenergetic beam from the photons that pass through.
Which of the following are true with regard to materials and attenuation? | Following the Compton effect, the wavelength change of the X-ray photon depends on the atomic number of the material | Photoelectric absorption is the predominating mode of interaction in modalities using high-energy photons | The Compton effect is the predominating mode of interaction in soft tissues | Photoelectric absorption is the predominating mode of interaction in air | The Compton effect is the predominating mode of interaction in contrast medium | FALSE | The change in a photon's wavelength after Compton scattering depends only on the scattering angle, not on the material it interacts with. | FALSE | In soft tissue, Compton scatter predominates at the medium-to-high energies used in CT and general radiography. The photoelectric effect dominates at lower energies. | TRUE | For most of the diagnostic X-ray energy range, Compton scatter is the most frequent interaction in low-Z materials like soft tissue, and it is the primary source of scatter. | FALSE | In low-Z materials like air, Compton scatter is the predominant interaction across most of the diagnostic energy range. | FALSE | In high-Z materials like iodine and barium contrast media, the photoelectric effect is the predominant interaction, which is why they produce such high image contrast.
Regarding secondary electrons: | Positrons are negatively charged electrons that result from radioactive decay | Beta particles can ionize atoms | The collision of two positrons results in two gamma photons, each with 511 keV | The range of the secondary electron is inversely proportional to the material density | Secondary electrons result in biological damage of tissues | FALSE | Positrons are the antimatter counterpart of electrons; they have a positive charge (+1). | TRUE | Beta particles are high-energy electrons or positrons. As they travel through matter, they interact with and ionize atoms along their path. | FALSE | The annihilation event occurs between a positron and an electron (matter and antimatter), not two positrons. | TRUE | An electron traveling through a denser material will undergo more frequent interactions, lose its energy more quickly, and thus travel a shorter distance (have a shorter range). | TRUE | The majority of biological damage from diagnostic X-rays is caused indirectly by the secondary electrons (photoelectrons and Compton electrons) that are set in motion by the initial photon interaction.
Which of the following are true with regard to filtration? | The aim of filtration is to make the beam intensity more uniform by removing the very high-energy rays | The tube housing acts as a valuable filter | The predominant attenuation process in a filter should be photoelectric absorption | At 80 kV, the half-value layer (HVL) of a beam with 2.5 mm Al filtration is typically measured as 2–3 mm Al | The thickness of copper required to reduce the intensity of an X-ray beam by a factor of 2 is greater than the required thickness of aluminium needed to have the same effect | FALSE | Filtration aims to remove low-energy ("soft") X-rays, which would increase patient skin dose without contributing to the diagnostic image. | TRUE | The X-ray beam is filtered by the glass tube, insulating oil, and metal housing before it exits the tube. This is known as the inherent filtration. | TRUE | Filters, typically made of aluminum, are designed to effectively absorb low-energy photons. In this energy range, the photoelectric effect is the dominant absorption mechanism. | TRUE | This is a standard quality control and regulatory benchmark. For a tube operating at 80 kVp, the total filtration must be sufficient to produce an HVL of at least 2.5 mm of aluminum. | FALSE | Copper (Z=29) is a much more effective attenuator than aluminum (Z=13). Therefore, a much thinner layer of copper is needed to achieve the same reduction in intensity.
Concerning the effects of filtration: | It increases the intensity of the beam | It increases the half-value layer (HVL) of the beam | It reduces the peak photon energy | It reduces the effective photon energy | It increases skin exit/entry dose ratio | FALSE | Filtration removes photons from the beam, so it always decreases the beam intensity (quantity). | TRUE | By removing the lower-energy photons, filtration "hardens" the beam, increasing its average energy and penetrating power, which results in a higher HVL. | FALSE | The peak photon energy is determined solely by the tube voltage (kVp) and is not affected by filtration. | FALSE | By selectively removing low-energy photons, filtration increases the average or effective energy of the remaining beam. | TRUE | A more penetrating (harder) beam passes through the patient more efficiently. This increases the dose at the exit surface relative to the entrance surface dose.
Decreasing the tube voltage (kV) results in which of the following? | It reduces the number of electrons colliding with the target | It reduces the photon fluence | It reduces the beam intensity | It results in the increase of the half-value layer (HVL) for a given material | The rate of photoelectric attenuation increases more than the rate of Compton attenuation | FALSE | The number of electrons (tube current) is controlled by the mA setting, which is independent of the kV. | TRUE | The efficiency of X-ray production decreases significantly at lower kV settings, resulting in fewer photons (lower fluence) for a given mAs. | TRUE | Both the quantity and the maximum energy of photons are reduced, leading to a substantial decrease in overall beam intensity. | FALSE | A lower kV produces a less penetrating beam with a lower average energy. This beam is more easily attenuated, so it decreases the HVL. | TRUE | The probability of the photoelectric effect is highly sensitive to energy (proportional to 1/E³), while the Compton effect is less so. As energy decreases, the photoelectric effect rapidly becomes more dominant.
Which of the following increase X-ray beam intensity at a given point from a source? | Increasing the filter thickness | Increasing the tube current | Increasing the distance | Increasing the tube voltage | Reducing the atomic number of the target material | FALSE | Increasing filtration removes photons from the beam and therefore decreases its intensity. | TRUE | Beam intensity is directly proportional to the tube current (mA). Doubling the mA doubles the number of photons produced per second. | FALSE | According to the inverse square law, increasing the distance from the source decreases the intensity. | TRUE | Increasing the kVp increases both the efficiency of X-ray production and the energy of the photons, resulting in a significant increase in beam intensity. | FALSE | The efficiency of Bremsstrahlung production is proportional to the atomic number (Z) of the target. Reducing Z decreases the X-ray output.
Regarding the atom, which of the following are true? | The mass number is equal to the number of nucleons | Positrons are found within the nucleus | The inner shell influences the chemical properties of an atom | The outermost shell with electrons is known as the valence shell | Radioactivity is dependent on the nucleus | TRUE | The mass number (A) is defined as the total count of protons and neutrons in the nucleus. Protons and neutrons are collectively called nucleons. | FALSE | The nucleus contains protons and neutrons. Positrons are antimatter particles emitted from the nucleus during certain types of radioactive decay. | FALSE | An atom's chemical properties are determined by the electrons in its outermost shell, known as the valence electrons. | TRUE | This is the definition of the valence shell. Its electrons are involved in forming chemical bonds. | TRUE | Radioactivity is the process by which an unstable atomic nucleus loses energy by emitting radiation. It is a purely nuclear phenomenon.
Which of the following are correct with regard to X-ray production? | Increasing the tube voltage increases the heat production at the target | Reducing the actual focal spot reduces the heat load on the target | Increasing the target angle increases the target heat rating for a given effective spot size | Increasing the target angle increases the effective focal spot | A rotating anode can take a higher heat load than a stationary anode | TRUE | Increasing the kV means more energy is delivered by each electron. Since >99% of this energy becomes heat, the total heat produced increases. | FALSE | Reducing the actual focal spot concentrates the heat into a smaller area, increasing the thermal stress and making it more difficult to dissipate heat. | TRUE | A larger (steeper) target angle spreads the electron beam over a larger physical area on the anode, improving heat dissipation for the same projected (effective) focal spot size. | TRUE | According to the line focus principle, for a given actual focal spot size, a larger target angle will produce a larger effective focal spot. | TRUE | By constantly presenting a new, cool surface to the electron beam, a rotating anode distributes the heat load over a much larger area, allowing it to withstand significantly higher exposures.
The nucleons in an atom: | Reduce following beta-positive (b+) decay | Are equal to the atomic number | Are equal to the difference between mass number and the number of protons | May have a negative charge | Affect the binding energy of electrons | FALSE | In beta-positive decay, a proton converts into a neutron. The total number of nucleons (protons + neutrons) remains unchanged. | FALSE | The atomic number (Z) is the number of protons only. Nucleons are the sum of protons and neutrons. | FALSE | The difference between the mass number (A) and the number of protons (Z) is the number of neutrons. The number of nucleons is equal to A. | FALSE | Nucleons are either positively charged (protons) or have no charge (neutrons). They cannot have a negative charge. | FALSE | The binding energy of electrons is determined by the electrostatic force from the positively charged protons. Neutrons have no charge and do not directly affect it.
Regarding photons: | X-ray photons are produced following K-electron capture | Collision of a positron with a negative beta particle results in the production of gamma photons | X-ray photons are produced following photoelectric absorption | They can be scattered only sideways or backwards following Compton attenuation | They have a mass | TRUE | When a nucleus captures a K-shell electron, it creates a vacancy. The filling of this vacancy by an outer-shell electron produces a characteristic X-ray. | TRUE | A negative beta particle is an electron. The annihilation of a positron (anti-electron) and an electron converts their mass into two 511 keV gamma photons. | TRUE | The photoelectric effect creates an inner-shell vacancy. When an outer-shell electron fills this hole, the energy difference is released as a characteristic X-ray photon. | FALSE | Compton scattered photons are predominantly scattered in the forward direction, especially at higher diagnostic energies. | FALSE | Photons are fundamental particles that are defined as having zero rest mass. They are packets of pure energy.
Regarding radioactivity: | Decay of radionuclides with a neutron excess produces a daughter nucleus with a higher atomic number | The rate of decay can be increased by heating the radionuclide | Isomers have the same half-life | Decay of radionuclides with a neutron deficit produces a daughter nucleus with a lower atomic number | Positron annihilation results in energy being converted to mass | TRUE | A neutron excess leads to beta-minus decay, where a neutron becomes a proton (n → p + e⁻). This increases the proton count, so the atomic number (Z) increases by one. | FALSE | Radioactive decay is a nuclear process whose rate is constant and unaffected by external physical or chemical conditions like temperature, pressure, or chemical state. | FALSE | Isomers are nuclides with the same Z and A but in different nuclear energy states. These different states (e.g., Tc-99 vs Tc-99m) have different half-lives. | TRUE | A neutron deficit leads to positron emission (p → n + e⁺) or electron capture (p + e⁻ → n). In both cases, a proton is converted to a neutron, so the atomic number (Z) decreases by one. | FALSE | Positron annihilation is a perfect example of mass being converted into energy according to Einstein's equation, E=mc².
Regarding the evolution of CT scanner generations: | First-generation scanners utilized a wide fan beam of X-rays and a large detector array | Third-generation scanners are characterized by a rotating fan beam and a detector arc that rotate together around the patient | Fourth-generation scanners employ a stationary X-ray tube and a rotating ring of detectors | The primary advantage of moving from first to third-generation technology was a significant improvement in spatial resolution | All modern multidetector CT (MDCT) scanners are based on a second-generation (rotate-translate) design | FALSE | First-generation scanners used a single, thin "pencil beam" and one or two detectors that performed a rotate-translate motion. | TRUE | This "rotate-rotate" geometry is the basis for virtually all modern CT scanners, where the X-ray tube and detector array are fixed relative to each other and rotate as a single unit. | FALSE | Fourth-generation scanners have a stationary ring of detectors, and only the X-ray tube rotates inside the ring. | FALSE | The primary advantage was a massive reduction in scan time, from minutes per slice to seconds per slice, making clinical CT practical. | FALSE | Modern MDCT scanners are an advanced evolution of the third-generation "rotate-rotate" design.
Concerning the principles of helical (spiral) CT: | Helical scanning requires the patient table to remain stationary while the gantry rotates continuously | Slip-ring technology allows for a maximum gantry rotation of 360 degrees before the gantry must stop and reverse | A pitch greater than 1.0 means that the patient table moves a distance greater than the X-ray beam width in one rotation | The primary advantage of helical CT over axial CT is its inherently higher spatial resolution | Data from a helical scan can be directly reconstructed into images without any mathematical correction | FALSE | Helical scanning is defined by the continuous movement of the patient table through the gantry as the gantry rotates. | FALSE | Slip-ring technology is what enables continuous, unending gantry rotation in one direction, which is the key innovation that allows for helical data acquisition. | TRUE | Pitch is the ratio of table travel per rotation to the beam collimation. A pitch > 1 results in a "stretched" helix and reduces patient dose. | FALSE | The main advantages are faster acquisition, elimination of inter-scan delay, and the ability to generate overlapping volumetric reconstructions from a single dataset. | FALSE | Because the data is acquired in a helical path, interpolation algorithms are required to synthesize planar slice data before the final image can be reconstructed.
Regarding Hounsfield Units (HU) and the CT scale: | The Hounsfield scale is an absolute measure of a material's physical density in g/cm³ | A Hounsfield Unit (HU) value of 0 corresponds to the radiodensity of air | Dense cortical bone typically has a HU value in the range of +40 to +100 HU | A negative HU value indicates that the material is less attenuating to X-rays than pure water | The HU value for a given tissue is independent of the X-ray beam energy (kVp) used for the scan | FALSE | The Hounsfield scale is a relative scale that linearizes the linear attenuation coefficient of a material relative to that of water. | FALSE | By definition, 0 HU corresponds to the radiodensity of pure water. Air has a value of approximately -1000 HU. | FALSE | Dense cortical bone is highly attenuating and has HU values typically greater than +1000 HU. | TRUE | This is the basis of the scale. Materials that are less attenuating than water, such as fat (approx. -50 to -100 HU) and air (-1000 HU), have negative values. | FALSE | Because attenuation coefficients (especially the photoelectric component) are energy-dependent, the HU value for a tissue can change slightly with different kVp settings.
Pertaining to CT image reconstruction: | Filtered back-projection is a modern technique that has completely replaced iterative reconstruction | The process of convolution involves applying a mathematical filter, or kernel, to the raw projection data | A "soft" or "smooth" reconstruction kernel is used to enhance the visibility of fine details and sharp edges | Iterative reconstruction techniques generally result in higher image noise compared to filtered back-projection at the same dose | The raw data stored after a scan contains the Hounsfield Units for each pixel | FALSE | Filtered back-projection (FBP) was the historical standard. Modern scanners increasingly use iterative reconstruction, which is more computationally intensive but can reduce noise and artifacts. | TRUE | This is a key step in filtered back-projection. Applying a kernel in the frequency domain corrects for the blurring that would result from simple back-projection alone. | FALSE | A "soft" or "smooth" kernel reduces noise and improves low-contrast resolution but at the expense of blurring fine details. A "sharp" or "bone" kernel enhances edges. | FALSE | A major advantage of iterative reconstruction is its ability to produce images with lower noise and/or allow for a significant reduction in patient dose compared to FBP. | FALSE | The raw data consists of the attenuation measurements (projections) from the detectors. This data must be processed through reconstruction algorithms to create the final image made of pixels with HU values.
Regarding key CT acquisition parameters: | Increasing the tube current (mA) primarily affects the spatial resolution of the image | Slice thickness is determined by the physical size of the individual detector elements in an MDCT scanner | A shorter gantry rotation time improves temporal resolution and reduces motion artifacts | A higher pitch value results in an increased radiation dose to the patient | The kVp setting controls the number of X-ray photons produced, while mAs controls their energy | FALSE | Increasing the mA increases the number of X-ray photons, which reduces image noise and improves contrast resolution. It does not directly affect spatial resolution. | FALSE | In MDCT, slice thickness is determined by combining the signals from a selected number of detector rows, not by the size of a single detector. | TRUE | A faster rotation allows the data for an image to be acquired more quickly, which "freezes" motion from things like the heartbeat, thereby improving temporal resolution. | FALSE | A higher pitch (e.g., 1.5) means the table travels faster for a given beam width, "stretching" the radiation dose over a longer volume. This reduces the patient dose. | FALSE | This is reversed. The kVp controls the peak energy (and quality) of the photons, while the mAs (mA multiplied by time) controls the total number (quantity) of photons produced.
Concerning CT image quality metrics: | Spatial resolution in CT is primarily limited by the patient's motion during the scan | Contrast resolution refers to the ability to distinguish between objects with very similar densities | Temporal resolution is a measure of the smallest object that can be resolved in an image | Image noise increases as the radiation dose (mAs) is increased | A large matrix size (e.g., 1024x1024) will result in lower spatial resolution than a small matrix size (512x512) for the same field of view | FALSE | Spatial resolution is limited by factors like focal spot size, detector aperture width, and reconstruction kernel. Motion causes artifacts but doesn't define the intrinsic resolution limit. | TRUE | This is the definition of contrast resolution, or low-contrast detectability. It is a key advantage of CT over radiography and is highly dependent on noise. | FALSE | Temporal resolution is the ability to resolve moving objects and is determined by the speed of data acquisition (e.g., gantry rotation time). | FALSE | Image noise is inversely related to the number of photons used. Therefore, noise decreases as the radiation dose (mAs) is increased. | FALSE | For a constant field of view, a larger matrix means smaller individual pixels. Smaller pixels allow for the resolution of finer details, resulting in higher spatial resolution.
On the use of windowing for CT image display: | The Window Width (WW) controls the brightness of the displayed image | The Window Level (WL) determines the range of Hounsfield Units that will be displayed as shades of gray | A narrow window width (e.g., 100 HU) is used to display tissues with a wide range of densities, like the entire abdomen | To view lung parenchyma, a window level of approximately +40 HU should be used | Setting a Window Level of 0 HU and a Window Width of 400 HU is a common setting for viewing soft tissues | FALSE | The Window Level (WL) controls the midpoint of the displayed HU range and thus determines the overall brightness of the image. | FALSE | The Window Width (WW) determines the range of HU values that are mapped to the available gray scale. The Window Level (WL) sets the center of that range. | FALSE | A narrow window width provides high contrast over a small range of densities. It is used to view tissues with similar densities, like the brain's gray and white matter. | FALSE | Lung parenchyma is mostly air and has a very low HU value. A window level of approximately -500 to -700 HU is used to properly visualize it. | TRUE | This "soft tissue" window (WL/WW of ~40/400) effectively displays the range of densities found in most abdominal and pelvic organs.
Regarding Multidetector CT (MDCT) technology: | MDCT scanners use a single row of very long detectors to capture the X-ray fan beam | The number of slices acquired per rotation is determined by the gantry rotation speed | Isotropic imaging means that the voxels are shaped like elongated cubes | The main advantage of MDCT is the ability to acquire volumetric data with thin slices in a short amount of time | In a 64-slice scanner, the X-ray beam is collimated to the width of a single detector row | FALSE | MDCT scanners are defined by having multiple parallel rows of detector elements along the patient's long axis (the z-axis). | FALSE | The number of slices acquired per gantry rotation is determined by the number of detector rows and how their data is combined by the data acquisition system (DAS). | FALSE | Isotropic imaging is a major goal of MDCT, where the voxels are perfect cubes (equal dimensions in x, y, and z), allowing for high-quality multiplanar reformats. | TRUE | This is the key benefit. MDCT allows for rapid acquisition of large volumes of data with very high z-axis resolution, enabling applications like CT angiography and detailed 3D rendering. | FALSE | In a 64-slice scanner, the beam is collimated to cover the entire width of the 64 active detector rows.
Pertaining to CT radiation dose: | The CT Dose Index (CTDI) is a direct measure of the total radiation dose received by a specific patient | The Dose-Length Product (DLP) is calculated by multiplying the CTDIvol by the patient's weight | Using automated tube current modulation can help reduce the overall patient dose | A smaller patient will always receive a lower absorbed dose than a larger patient for the same CT protocol | The effective dose from a CT scan can be estimated by multiplying the DLP by a region-specific conversion factor | FALSE | The CTDI is a standardized measure of dose delivered to a reference phantom (either 16 cm or 32 cm). It is an indicator of scanner output, not the dose to an individual patient. | FALSE | The DLP is calculated by multiplying the volumetric CTDI (CTDIvol) by the scan length in centimeters. It represents the total energy imparted over the entire scan volume. | TRUE | Automated tube current modulation adjusts the mA based on the patient's attenuation at different locations (e.g., lower mA at the lungs, higher at the shoulders), optimizing image quality and reducing dose. | FALSE | For the same technique factors, a smaller patient will have less tissue to attenuate the beam, resulting in a higher absorbed dose. Protocols must be adjusted for patient size. | TRUE | This is the standard method for estimating the overall cancer risk. Conversion factors (k-factors) exist for different body regions (head, chest, abdomen) to convert DLP into an effective dose (mSv).
Concerning common CT artifacts: | Beam hardening artifact causes dark bands or streaks to appear between two dense objects | Partial volume averaging occurs when a voxel contains tissues with very similar Hounsfield Units | Motion artifact primarily manifests as blurring and ghosting in the direction of gantry rotation | Metal artifact is caused by the metal absorbing all the X-rays, leaving a signal void on the image | A ring artifact is characteristic of a malfunction in a third-generation CT scanner | TRUE | This occurs because the X-ray beam becomes "harder" (higher average energy) as it passes through the first dense object, causing it to be less attenuated than expected by the second object. | FALSE | Partial volume averaging occurs when a single voxel contains tissues with different Hounsfield Units (e.g., bone and brain). The scanner assigns an average HU value, obscuring detail. | FALSE | While motion causes blurring, the characteristic ghosting artifacts appear in the direction that takes the longest to acquire, which in CT corresponds to the direction of patient motion through the gantry. | FALSE | Metal artifact is caused by a combination of severe beam hardening and photon starvation, which leads to characteristic bright and dark streaking artifacts radiating from the metal object. | TRUE | Ring artifacts are caused by a miscalibrated or faulty detector element. In a third-generation (rotate-rotate) design, this faulty detector traces a circular path in the data, resulting in a ring on the image.
Regarding Dual-Energy CT (DECT): | DECT systems use two separate X-ray tubes that are operated at the same kVp | The principle of DECT relies on the fact that all materials have the same attenuation coefficient at different energies | DECT allows for the creation of "virtual non-contrast" images from a contrast-enhanced scan | A primary application of DECT is to reduce motion artifacts from breathing | DECT is unable to differentiate between different types of kidney stones | FALSE | DECT systems acquire data at two different energy levels, either by using two tubes at different kVp settings (e.g., 80 kVp and 140 kVp) or by rapidly switching the kVp of a single tube. | FALSE | The principle relies on the fact that different materials (e.g., iodine vs. calcium) have different changes in their attenuation coefficients at different X-ray energies. | TRUE | By identifying materials based on their energy-dependent attenuation, DECT can create a virtual non-contrast image by computationally removing the signal from iodine contrast material. | FALSE | While DECT has many applications, reducing motion artifacts is not one of them. Its strengths are in material decomposition, artifact reduction (metal), and tissue characterization. | FALSE | A major application of DECT is differentiating between uric acid stones and calcium-based stones, as they have different attenuation properties at different energies, which aids in treatment planning.
Concerning CT Angiography (CTA): | CTA requires the use of a negative oral contrast agent to visualize blood vessels | For optimal arterial enhancement, the scan should be initiated immediately after the injection of intravenous contrast | Bolus tracking is a technique where the scanner monitors the patient's heart rate to time the scan | A test bolus injection can be used to determine the patient-specific time it takes for contrast to reach the target artery | CTA provides better spatial resolution for vessel imaging than conventional digital subtraction angiography (DSA) | FALSE | CTA relies on the injection of a high-concentration intravenous iodinated contrast agent to make the blood vessels appear bright (hyperdense). | FALSE | There is a delay between injection and arterial arrival. Scanning too early will miss the peak enhancement. The timing must be precise to capture the arterial phase. | FALSE | Bolus tracking involves placing a region of interest (ROI) over a target vessel (e.g., the aorta) and automatically triggering the scan once the contrast enhancement reaches a predefined HU threshold. | TRUE | This is a common method. A small amount of contrast is injected and timed, allowing for precise calculation of the delay needed for the main diagnostic scan. | FALSE | While CTA offers excellent 3D visualization, conventional DSA (an interventional radiology technique) still has superior spatial resolution for visualizing very small vessels.
Regarding Cone-Beam CT (CBCT): | CBCT uses a narrow, fan-shaped X-ray beam similar to conventional MDCT | The detector used in CBCT is typically a curved array of scintillation crystals | A complete volumetric dataset in CBCT is acquired during a single 360-degree rotation of the gantry | CBCT provides superior low-contrast resolution (soft-tissue contrast) compared to MDCT | A major advantage of CBCT is its significantly lower radiation dose compared to a similar MDCT scan | FALSE | As the name implies, CBCT uses a large, cone-shaped X-ray beam to acquire a volumetric dataset of the entire field of view at once. | FALSE | The detector in CBCT is typically a large, flat-panel amorphous silicon or CMOS detector, similar to those used in digital radiography. | TRUE | This is a key feature. The cone beam covers the entire volume of interest, allowing all the necessary data to be acquired in a single rotation. | FALSE | Due to a higher amount of detected scatter and limitations of flat-panel detectors, CBCT has significantly poorer low-contrast resolution than conventional MDCT. | TRUE | Because it is often used for bone and high-contrast imaging and requires only a single rotation, the radiation dose from a CBCT scan is generally much lower than for a diagnostic MDCT study of the same area.
On the topic of CT image noise and artifacts: | Increasing the slice thickness will increase the amount of image noise | The "windmill" artifact is a type of motion artifact seen in helical cardiac CT | Partial volume averaging can be reduced by using thicker reconstruction slices | Ring artifacts are most commonly seen on fourth-generation (rotate-stationary) CT scanners | Photon starvation artifacts are most likely to occur when scanning through the shoulders or pelvis | FALSE | Increasing the slice thickness includes more photons per voxel, which improves the signal-to-noise ratio and thus decreases image noise. | TRUE | This artifact is caused by the helical acquisition path interacting with the periodic motion of the heart, especially when the heart rate is not stable or synchronized with the scan. | FALSE | Partial volume averaging occurs when different tissues are averaged within a single voxel. It is reduced by using thinner slices, which results in smaller voxels. | FALSE | Ring artifacts are caused by a faulty detector in a third-generation (rotate-rotate) scanner, where the detector's path is circular relative to the isocenter. | TRUE | Photon starvation occurs when the X-ray beam is so heavily attenuated by dense body parts (like the shoulders) that too few photons reach the detectors, leading to severe streaking and noise.
Regarding CT scanner hardware and technology: | The X-ray tube in a modern CT scanner is typically a stationary anode design | Bowtie filters are used to absorb photons at the periphery of the beam, creating a more uniform signal at the detector | The Data Acquisition System (DAS) is located within the X-ray tube housing | The gantry aperture refers to the maximum scan field of view that can be reconstructed | Slip rings transmit power and data to the rotating gantry components using physical wires that must be rewound after each scan | FALSE | Due to the immense heat loads, CT scanners use high-performance rotating anode X-ray tubes. | TRUE | Patients are generally cylindrical, so the beam path is shorter at the edges than in the middle. The bowtie filter selectively attenuates the periphery to reduce patient skin dose and prevent detector overload. | FALSE | The DAS is located in the rotating gantry, near the detectors. Its function is to measure the detector signals, convert them to digital format, and transmit them for reconstruction. | FALSE | The gantry aperture is the physical opening, or bore, in the scanner through which the patient passes (typically 70-80 cm). The scan field of view is usually smaller (~50 cm). | FALSE | Slip rings use a system of brushes and conductive rings to allow wireless transmission of power and data, enabling the continuous, unlimited rotation essential for helical scanning.
Regarding radioactive decay and half-life: | Alpha decay involves the emission of a high-energy electron from the nucleus | Beta-minus decay occurs in neutron-rich nuclides and results in an increase in the atomic number (Z) | The effective half-life of a radiopharmaceutical is always longer than its physical half-life | Electron capture and positron emission are competing decay processes that occur in proton-rich nuclides | A gamma ray is a charged particle emitted during the de-excitation of a nucleus | FALSE | Alpha decay is the emission of an alpha particle, which is a helium nucleus consisting of 2 protons and 2 neutrons. | TRUE | In beta-minus decay, a neutron transforms into a proton (n → p + e⁻), thus increasing the atomic number by one. | FALSE | The effective half-life combines physical decay and biological clearance and is always shorter than the shorter of the two individual half-lives. | TRUE | Both processes effectively convert a proton into a neutron to stabilize a proton-rich (neutron-deficient) nucleus, causing the atomic number (Z) to decrease by one. | FALSE | A gamma ray is an uncharged, high-energy photon of electromagnetic radiation.
Pertaining to Technetium-99m ($^{99m}$Tc): | Technetium-99m is directly produced for clinical use inside a nuclear reactor | An ideal diagnostic radionuclide should primarily emit particulate radiation like alpha or beta particles | The 140 keV gamma emission of $^{99m}$Tc is well-suited for imaging with modern gamma cameras | The term "elution" refers to the process of injecting the radiopharmaceutical into the patient | The physical half-life of $^{99m}$Tc is approximately 67 hours | FALSE | The parent isotope, Molybdenum-99 (⁹⁹Mo), is produced in a reactor. $^{99m}$Tc is then separated from ⁹⁹Mo in a radionuclide generator on-site at the hospital. | FALSE | An ideal radionuclide for imaging should emit only gamma rays of a single energy to maximize image quality and minimize patient radiation dose from non-penetrating particles. | TRUE | This energy is high enough to escape the patient's body but low enough to be efficiently detected by the NaI(Tl) crystal and effectively collimated by lead collimators. | FALSE | "Elution" is the process of washing the daughter radionuclide ($^{99m}$Tc) off the generator column using saline. This is how the radiopharmaceutical is prepared. | FALSE | The physical half-life of $^{99m}$Tc is approximately 6 hours. Its parent, ⁹⁹Mo, has a half-life of about 66-67 hours.
Regarding gamma camera collimators: | A parallel-hole collimator is designed to magnify the image of a small organ | Pinhole collimators provide the highest sensitivity (count rate) of all collimator types | Converging-hole collimators are used to make a large organ, like the liver, appear smaller to fit in the field of view | The primary function of a collimator is to absorb gamma rays that are not traveling perpendicular to the detector crystal | The thickness of the septa (walls) in a collimator primarily determines its spatial resolution | FALSE | A parallel-hole collimator projects an image that is the same size as the organ, with no magnification or minification. | FALSE | Pinhole collimators have very low sensitivity because their small aperture blocks the vast majority of photons. They offer excellent spatial resolution and magnification. | FALSE | Converging-hole collimators magnify the image, making small organs (like the thyroid) appear larger on the detector face. | TRUE | The lead septa of the collimator absorb photons traveling at an oblique angle, ensuring that only photons traveling parallel to the holes are detected, which allows for spatial localization. | FALSE | The thickness of the septa determines the maximum photon energy the collimator can handle without "septal penetration." The length and diameter of the holes primarily determine the resolution/sensitivity trade-off.
Pertaining to gamma camera detector systems: | The scintillation crystal in a modern gamma camera is typically made of Bismuth Germanate (BGO) | A photomultiplier tube (PMT) converts the energy of a gamma ray directly into an electrical signal | A larger number of smaller PMTs results in improved spatial resolution | The output signal from a single PMT is sufficient to determine the exact location of a scintillation event | Pulse Height Analysis (PHA) is used to reject photons that have undergone Compton scatter within the patient | FALSE | The standard scintillation crystal used in most gamma cameras is Thallium-activated Sodium Iodide (NaI(Tl)) due to its high light output and good energy resolution. | FALSE | A PMT detects the very weak light photons produced by the scintillation crystal and converts them into a much larger, measurable electrical signal through amplification. | TRUE | Using more, smaller PMTs allows the positioning logic circuit to more accurately determine the centroid of the light spread from a scintillation event, thus improving intrinsic spatial resolution. | FALSE | The location is determined by a weighted average of the signals from all PMTs that detect light from a single event. The PMT closest to the event will have the strongest signal. | TRUE | Scattered photons have lower energy than primary photons. The PHA window is set around the photopeak energy (e.g., 140 keV for $^{99m}$Tc) to accept only unscattered photons, improving image contrast.
Regarding SPECT (Single Photon Emission Computed Tomography): | SPECT imaging provides only two-dimensional, planar views of radionuclide distribution | The gamma camera gantry remains stationary during a SPECT acquisition | Filtered back-projection in SPECT reconstruction does not require any correction for attenuation | Attenuation correction in SPECT is less critical than in PET because the photons have lower energy | The primary advantage of SPECT over planar imaging is its ability to quantify absolute tracer concentration | FALSE | SPECT is a tomographic technique that acquires data from multiple angles around the patient to reconstruct cross-sectional (3D) images, similar to CT. | FALSE | For SPECT, the gamma camera head(s) rotate around the patient, acquiring a series of planar images (projections) from many different angles. | FALSE | Attenuation (the absorption and scatter of photons within the patient) is a significant issue in SPECT. Attenuation correction is crucial for accurate image reconstruction. | FALSE | Attenuation correction is more complex and arguably more critical in SPECT because the photons have lower energy (e.g., 140 keV) and are more subject to attenuation than the 511 keV photons in PET. | FALSE | While quantitative SPECT is possible, its primary clinical advantage over planar imaging is the removal of overlapping structures and improved image contrast and lesion localization.
Concerning PET (Positron Emission Tomography): | The fundamental event detected in PET is the direct interaction of a positron with the detector | The two gamma photons produced during annihilation are emitted at a 90-degree angle to each other | For an annihilation event to be recorded, the two photons must strike opposing detectors simultaneously | The energy of each annihilation photon is 140 keV | "Time-of-Flight" (TOF) PET uses the decay time of the radionuclide to improve image quality | FALSE | The fundamental event is the coincidence detection of the pair of 511 keV gamma photons that result from the annihilation of a positron with an electron. | FALSE | To conserve momentum, the two annihilation photons are emitted in opposite directions, at an angle of approximately 180 degrees to each other. | TRUE | This is the principle of "coincidence detection." The electronics look for two signals arriving within a very short time window (a few nanoseconds) along a line of response (LOR). | FALSE | The energy of each of the two annihilation photons is 511 keV, which corresponds to the rest mass energy of an electron (or positron). | FALSE | TOF PET uses the slight difference in the arrival time of the two photons at the detectors to more precisely localize the annihilation event along the line of response, improving image quality.
Regarding PET/CT and PET/MRI hybrid imaging: | The CT component of a PET/CT scanner is primarily used for radiation therapy planning | In PET/CT, the CT scan is used to generate an attenuation correction map for the PET data | PET/MRI is not clinically feasible because the MRI's magnetic field interferes with the PET detectors | The primary advantage of PET/MRI over PET/CT is its significantly faster scan time | The radiation dose to the patient from a PET/MRI scan is much higher than from a PET/CT scan | FALSE | While the fused images are often used for therapy planning, the primary function of the CT component during the scan is for anatomical localization and attenuation correction. | TRUE | The CT Hounsfield Units can be converted into a map of linear attenuation coefficients for the 511 keV PET photons, providing a patient-specific, high-quality attenuation correction. | FALSE | PET/MRI is clinically feasible. It requires specialized PET detectors (e.g., using silicon photomultipliers or avalanche photodiodes) that are not affected by strong magnetic fields. | FALSE | A major challenge of PET/MRI is its longer scan time compared to the very rapid acquisition of a PET/CT scan. | FALSE | Because MRI does not use ionizing radiation, the total radiation dose to the patient from a PET/MRI scan is significantly lower than from a PET/CT scan.
Pertaining to image quality in Nuclear Medicine: | Increasing the acquisition time for a scan will decrease the image noise | Spatial resolution in nuclear medicine imaging is superior to that of CT or MRI | The sensitivity of a gamma camera is improved by using a high-resolution collimator | A "star artifact" is typically caused by patient motion during a SPECT scan | Intrinsic uniformity refers to the performance of the collimator-detector system together | TRUE | Image noise is dominated by statistical fluctuations in the number of detected counts. Acquiring for a longer time increases the total counts, which improves the signal-to-noise ratio. | FALSE | Nuclear medicine imaging has the poorest spatial resolution of all major imaging modalities, typically on the order of 5-10 mm, compared to sub-millimeter resolution for CT and MRI. | FALSE | A high-resolution collimator has smaller holes and thicker septa, which improves resolution but decreases sensitivity because it accepts fewer photons. | FALSE | A "star artifact" is a reconstruction artifact that appears as streaks radiating from a very "hot" (high-activity) area in the image. | FALSE | Intrinsic uniformity is a quality control test of the detector crystal and electronics without the collimator attached. Extrinsic uniformity tests the entire system with the collimator.
Regarding Radionuclide Therapy: | Therapeutic radionuclides are chosen because they emit highly penetrating gamma rays | Iodine-131 is used to treat thyroid cancer because it is a pure alpha emitter | The goal of radionuclide therapy is to deliver a diagnostic-level dose to diseased tissue | A "theranostic" agent is one that can be used for both diagnostic imaging and targeted therapy | Beta particles are ideal for therapy because their long range in tissue allows them to treat large, diffuse tumors | FALSE | Therapeutic radionuclides are chosen because they emit short-range, highly damaging particulate radiation (e.g., beta or alpha particles) that deposits energy locally in the target tissue. | FALSE | Iodine-131 is used for thyroid therapy because it is primarily a beta emitter, which destroys the thyroid cells. It also emits gamma rays that can be imaged. | FALSE | The goal is to deliver a very high, ablative (cytocidal) radiation dose to the target tissue while sparing surrounding healthy organs. | TRUE | This is the definition of theranostics (therapy + diagnostics), where the same molecular target is used for both imaging (with a gamma or positron emitter) and treatment (with a beta or alpha emitter). | FALSE | Beta particles are ideal because their short range (typically a few millimeters) confines the radiation dose to the tumor and immediately surrounding tissue, sparing distant healthy organs.
Pertaining to internal dosimetry: | The MIRD methodology stands for "Medical Internal Radiation Dose" | The "source organ" is the organ that is most sensitive to radiation damage | The "absorbed fraction" is the fraction of activity administered that is taken up by an organ | The biological half-life of a radiopharmaceutical is a constant physical property of the radionuclide | The cumulative activity is independent of the physical half-life of the radionuclide | TRUE | The MIRD Committee of the Society of Nuclear Medicine developed the standard schema for calculating the radiation dose absorbed by a patient from an internal radiopharmaceutical. | FALSE | The "source organ" is the organ that contains the radionuclide (i.e., the source of the radiation). The "target organ" is the organ for which the dose is being calculated. | FALSE | This describes the uptake fraction. The "absorbed fraction" is the fraction of energy emitted by the source organ that is absorbed by the target organ. | FALSE | The biological half-life describes the rate of clearance of the pharmaceutical from the body and is dependent on the patient's physiology, not the radionuclide itself. | FALSE | Cumulative activity, which is proportional to the total number of decays, is highly dependent on both the administered activity and the effective half-life (which includes the physical half-life).
Regarding Quality Control (QC) in Nuclear Medicine: | A daily "flood" or uniformity test is performed to check the camera's spatial resolution | A bar phantom is used to assess the energy resolution of the gamma camera | The "Center of Rotation" (COR) test is a critical QC procedure for PET scanners | An extrinsic flood is performed with the collimator removed from the camera head | A "hot" or "cold" spot on a uniformity flood image may indicate a defective photomultiplier tube (PMT) | FALSE | A daily flood test is performed to check for non-uniformities in the detector's response, which ensures that a uniform source of activity produces a uniform image. | FALSE | A bar phantom, which consists of lead bars of varying widths, is used to measure the spatial resolution and linearity of the system. | FALSE | COR error is a major source of artifacts in SPECT imaging. This test ensures that the electronic center of the matrix aligns with the mechanical axis of rotation. | FALSE | An intrinsic flood is performed without the collimator. An extrinsic flood is performed with the collimator attached to assess the entire system performance. | TRUE | A defective PMT will result in a mispositioning of events in its vicinity, leading to a visible cold (or sometimes hot) spot on the flood image, indicating the need for service.
Pertaining to PET artifacts and corrections: | Attenuation in PET is less of a problem than in SPECT because the photons have higher energy | Random coincidences occur when two photons from two different annihilation events are detected simultaneously | Scatter correction is not necessary in PET because of the use of electronic collimation | A "cold" artifact can be produced by patient motion during the emission scan | The CT-based attenuation correction can cause artifacts if there is metal present in the patient | FALSE | While 511 keV photons are more penetrating, the total path length is longer (involving two photons traveling through the whole body width), making attenuation a very significant problem that requires correction. | TRUE | These are a major source of background noise in PET images. They are unrelated, uncorrelated events that happen to fall within the coincidence timing window by chance. | FALSE | Although there is no physical collimator, photons can still undergo Compton scatter in the patient and be detected, degrading image quality. Scatter correction is essential. | TRUE | If a patient moves during the long PET scan but not during the fast CT scan, the attenuation correction map will be misaligned with the emission data, leading to incorrect calculations and artifacts. | TRUE | The high Hounsfield units from metal on the CT scan will be misinterpreted as extremely dense tissue, leading the algorithm to over-correct for attenuation in that area, creating severe artifacts.
Regarding different nuclear medicine imaging modalities: | A bone scan is an example of a PET imaging procedure | Cardiac perfusion imaging can be performed using either SPECT or PET | Gallium-67 is a positron-emitting radionuclide used for tumor imaging | A V/Q scan is used to evaluate for deep vein thrombosis | FDG (Fluorodeoxyglucose) is a radiotracer that measures cellular glucose metabolism | FALSE | A bone scan is a classic planar or SPECT imaging procedure, most commonly performed with Technetium-99m labeled MDP. | TRUE | Myocardial perfusion can be assessed with SPECT tracers (like $^{99m}$Tc-sestamibi) or PET tracers (like Rubidium-82 or N-13 ammonia), with PET generally offering better image quality. | FALSE | Gallium-67 is a gamma-emitting radionuclide. Fluorine-18 (in FDG) is the most common positron emitter for tumor imaging. | FALSE | A V/Q (Ventilation/Perfusion) scan is a lung scan used to evaluate for pulmonary embolism. | TRUE | FDG is a glucose analog that is taken up by metabolically active cells. Its uptake, imaged with PET, is a surrogate for glucose metabolism, which is high in most cancers.
Pertaininng to scintillation detector properties: | A good scintillator should have a high light output (many light photons per keV of gamma energy) | "Hygroscopic" means that the detector material is resistant to absorbing water from the air | The decay time of a scintillator refers to its physical half-life | Energy resolution is the ability of the detector to distinguish between two separate points of light | NaI(Tl) crystals are ideal for Time-of-Flight PET because of their very fast light decay time | TRUE | High light output is crucial because it improves the signal for the PMTs, leading to better energy resolution and spatial resolution. | FALSE | "Hygroscopic" means the material readily absorbs moisture from the air. NaI(Tl) is highly hygroscopic and must be hermetically sealed in a casing to prevent damage. | FALSE | The decay time is the time it takes for the scintillation light flash to be emitted after a gamma ray interaction. A short decay time is crucial for high count-rate applications. | FALSE | This describes spatial resolution. Energy resolution is the ability of a detector to distinguish between gamma rays of slightly different energies. | FALSE | NaI(Tl) has a relatively slow decay time (~230 ns). Modern TOF PET scanners use very fast scintillators like LSO or LYSO, which have decay times of around 40 ns.
Regarding radiation safety in nuclear medicine: | After being injected with a radiopharmaceutical, a patient is not a source of radiation exposure to others | Spilling a therapeutic dose of Iodine-131 is considered a minor administrative issue | A "dose calibrator" is used to measure the radiation dose being received by a technologist | Most radiopharmaceuticals can be disposed of in regular trash after injection | Time, distance, and shielding are the primary principles for minimizing dose from patients containing radioactivity | FALSE | A patient injected with a radionuclide becomes a source of radiation (primarily gamma rays) and must be treated as such, with safety precautions taken by staff and family. | FALSE | Iodine-131 is a volatile therapeutic radionuclide. A spill constitutes a major radiation safety event that requires immediate, specific decontamination procedures. | FALSE | A dose calibrator is an ionization chamber used to accurately measure the activity (in Bq or Curies) of a radiopharmaceutical dose before it is administered to the patient. | FALSE | Radiopharmaceuticals must be stored for decay in a designated shielded area until their activity has fallen to background levels before they can be disposed of. | TRUE | These cardinal principles are just as important here as in X-ray imaging: minimize time spent near the patient, maximize distance, and use shielding (e.g., lead shields, leaded glass) when appropriate.
Regarding Multidetector CT (MDCT) technology: | MDCT scanners use a single row of very long detectors to capture the X-ray fan beam | The number of slices acquired per rotation is determined by the gantry rotation speed | Isotropic imaging means that the voxels are shaped like elongated cubes | The main advantage of MDCT is the ability to acquire volumetric data with thin slices in a short amount of time | In a 64-slice scanner, the X-ray beam is collimated to the width of a single detector row | FALSE | MDCT scanners are defined by having multiple parallel rows of detector elements along the patient's long axis (the z-axis). | FALSE | The number of slices acquired per gantry rotation is determined by the number of detector rows and how their data is combined by the data acquisition system (DAS). | FALSE | Isotropic imaging is a major goal of MDCT, where the voxels are perfect cubes (equal dimensions in x, y, and z), allowing for high-quality multiplanar reformats. | TRUE | This is the key benefit. MDCT allows for rapid acquisition of large volumes of data with very high z-axis resolution, enabling applications like CT angiography and detailed 3D rendering. | FALSE | In a 64-slice scanner, the beam is collimated to cover the entire width of the 64 active detector rows.
Regarding fundamental X-ray interactions and production: | In the diagnostic energy range, the probability of the photoelectric effect increases as photon energy increases. | Compton scatter is the primary contributor to image contrast when using iodinated contrast media. | The energy of characteristic radiation is dependent on the tube voltage (kVp). | Bremsstrahlung radiation produces a continuous spectrum of X-ray energies. | Beam hardening results in a decrease in the half-value layer (HVL) of an X-ray beam. | FALSE | The probability of the photoelectric effect is inversely proportional to approximately the cube of the photon energy (1/E³), meaning it decreases rapidly as photon energy increases. | FALSE | The high atomic number of iodine causes the photoelectric effect to be the dominant interaction, which is responsible for the high subject contrast seen with these agents. | FALSE | The energy of characteristic radiation is determined by the difference in electron binding energies of the target atom's shells, which is constant for a given material and independent of the kVp. | TRUE | Bremsstrahlung ("braking radiation") results from electrons decelerating at various rates, producing a continuous range of photon energies up to the maximum kVp. | FALSE | Beam hardening removes low-energy photons, increasing the average energy and penetrability of the beam. This results in an increase in the HVL.
Pertaining to digital radiography and image quality: | The dynamic range of a digital detector is narrower than that of a film-screen system. | Detective Quantum Efficiency (DQE) is a measure of the spatial resolution of an imaging system. | A digital imaging system with a higher DQE can produce the same image quality at a lower patient dose. | In Computed Radiography (CR), the latent image is read by exposing the imaging plate to an intense white light. | The DICOM standard ensures interoperability by dictating the specific exposure factors for all procedures. | FALSE | Digital detectors have a much wider, linear dynamic range, allowing them to accurately capture a broader range of exposure intensities without loss of information. | FALSE | DQE is a measure of the efficiency of a detector in converting incident X-ray photons into a useful image signal, directly relating to the signal-to-noise ratio performance. | TRUE | A higher DQE means the detector is more efficient at using the available photons. Therefore, fewer photons (and a lower dose) are required to achieve a given signal-to-noise ratio. | FALSE | The latent image in a CR plate is read by scanning it with a fine-point red laser, which stimulates the release of stored energy as blue-green light. | FALSE | DICOM is a standard for file formats and network communication to ensure that images and data can be shared and viewed between different systems; it does not dictate clinical protocols.
Concerning the principles of radiation safety and dosimetry: | Deterministic effects of radiation are characterized by a dose threshold below which they do not occur. | The risk of inducing a fatal cancer is an example of a deterministic effect. | The ALARA principle states that all radiation exposures must be kept below a legally defined dose limit. | The effective dose (Sv) is calculated by multiplying the absorbed dose (Gy) by a tissue weighting factor (wT​). | For the same absorbed dose, alpha particles have a higher equivalent dose than X-rays. | TRUE | This is the defining feature of deterministic effects (e.g., skin erythema, cataracts). The severity of the effect is also dose-dependent above the threshold. | FALSE | Cancer induction is the primary example of a stochastic effect, where the probability of occurrence, not the severity, is dependent on dose. | FALSE | ALARA (As Low As Reasonably Achievable) is a guiding principle to minimize dose at all times. Dose limits are the legal maximums that should not be exceeded. | FALSE | Effective dose is the sum of equivalent doses to each tissue multiplied by that tissue's weighting factor (wT​). Equivalent dose is absorbed dose multiplied by the radiation weighting factor (wR​). | TRUE | Alpha particles are densely ionizing and have a high radiation weighting factor (wR​=20), whereas X-rays have a wR​ of 1. Therefore, alpha particles have a much higher equivalent dose.
Regarding CT scanner technology and image formation: | The Hounsfield scale is an absolute measure of a material's physical density. | Modern multidetector CT (MDCT) scanners are based on a first-generation (rotate-translate) design. | A "soft" reconstruction kernel is used to reduce image noise at the expense of spatial resolution. | The use of a pitch greater than 1.0 in helical CT increases the radiation dose to the patient. | Windowing is a preprocessing technique automatically applied by the scanner to set image contrast. | FALSE | The Hounsfield scale is a relative scale that linearizes the linear attenuation coefficient of a material with respect to that of pure water (which is defined as 0 HU). | FALSE | Modern MDCT scanners are an evolution of the third-generation (rotate-rotate) design, where the X-ray tube and detector arc rotate together around the patient. | TRUE | A soft/smooth kernel averages pixel values, which reduces statistical noise and improves low-contrast visibility, but also blurs sharp edges and fine details, reducing spatial resolution. | FALSE | A pitch greater than 1.0 means the patient table travels further per gantry rotation, "stretching" the helical path. This reduces the patient radiation dose. | FALSE | Windowing is the primary post-processing tool used by operators to manually adjust the displayed brightness (window level) and contrast (window width) for optimal viewing.
Pertaining to fundamental MRI principles: | The Larmor frequency is directly proportional to the strength of the gradient fields. | T1 relaxation describes the recovery of longitudinal magnetization, while T2 relaxation describes the decay of transverse magnetization. | In biological tissues, the T1 relaxation time is always shorter than the T2 relaxation time. | A 90-degree radiofrequency (RF) pulse is used to convert longitudinal magnetization into transverse magnetization. | Gadolinium-based contrast agents work by lengthening the T1 relaxation time of adjacent tissues. | FALSE | The Larmor frequency is directly proportional to the strength of the main static magnetic field (B₀), not the much weaker, transient gradient fields. | TRUE | This is the correct definition of the two primary, independent relaxation processes that form the basis of MR image contrast. | FALSE | In biological tissues, T1 relaxation is a much slower process than T2 relaxation. Therefore, the T1 time constant is always significantly longer than the T2 time constant. | TRUE | At equilibrium, magnetization is aligned with B₀ (longitudinal). A 90-degree RF pulse applied at the Larmor frequency provides the energy to tip this magnetization into the transverse plane where it can be measured. | FALSE | Gadolinium is paramagnetic and shortens the T1 relaxation time of nearby water protons, causing tissues that take it up to appear much brighter on T1-weighted images.
Regarding Nuclear Medicine instrumentation and physics: | The primary function of a collimator in a gamma camera is to focus gamma rays onto the crystal. | The annihilation of a positron with an electron results in the emission of two 511 keV photons at a 90-degree angle to each other. | SPECT imaging creates tomographic slices by using a stationary detector and a rotating patient. | The physical half-life of Technetium-99m ($^{99m}$Tc) is approximately 6 hours. | A PET/CT scanner uses the CT component primarily for anatomical localization and attenuation correction of the PET data. | FALSE | Gamma rays cannot be focused with lenses. A collimator works by absorbing photons that are not traveling on a path perpendicular to the detector, thus allowing for spatial localization. | FALSE | To conserve momentum, the two annihilation photons are emitted in opposite directions, at approximately 180 degrees to each other. This geometry is fundamental to PET imaging. | FALSE | In SPECT, the patient remains stationary while the gamma camera gantry rotates around the patient to acquire projections from multiple angles. | TRUE | This relatively short half-life is ideal for diagnostic imaging, as it allows for sufficient imaging time while minimizing the patient's radiation dose. | TRUE | This is the key advantage of the hybrid system. The CT provides a detailed anatomical map and a patient-specific map of X-ray attenuation that is used to correct the PET emission data.
Concerning fluoroscopy and interventional radiology: | In an Image Intensifier, magnification mode is used to decrease the patient dose rate. | The Dose Area Product (DAP) is the best metric for predicting the risk of deterministic skin injury. | Digital Subtraction Angiography (DSA) works by subtracting a pre-contrast "mask" image from subsequent contrast-filled images. | The primary source of radiation exposure to staff in the interventional suite is leakage radiation from the X-ray tube. | Positioning the X-ray tube under the patient (under-couch) is preferred for minimizing operator dose. | FALSE | Magnification mode is less efficient, so the Automatic Brightness Control (ABC) system increases the tube output (mA/kVp) to maintain image brightness, which increases the patient dose rate. | FALSE | Cumulative Air Kerma at the interventional reference point is the best predictor of peak skin dose and deterministic injury risk. DAP is a better measure of stochastic risk. | TRUE | This subtraction process removes stationary background anatomy like bone, leaving a high-contrast image of only the contrast-filled blood vessels. | FALSE | The primary source of staff exposure is scatter radiation originating from the patient's body. | TRUE | An under-couch tube configuration directs the most intense scatter radiation towards the floor, away from the operator's head and torso, making it the safest setup.
Regarding Ultrasound Physics and Principles: | Ultrasound imaging uses high-frequency electromagnetic waves to generate images. | Acoustic impedance is the product of the tissue's density and the speed of sound within that tissue. | A large difference in acoustic impedance between two tissues results in strong reflection of the ultrasound beam at their interface. | The Doppler effect in ultrasound is used to measure the attenuation of sound in tissue. | The piezoelectric effect is the principle by which ultrasound probes convert electricity into sound and vice-versa. | FALSE | Ultrasound uses high-frequency mechanical sound waves (typically 2-15 MHz), not electromagnetic radiation. | TRUE | This is the correct definition. Acoustic impedance (Z) is a fundamental property of a tissue that determines how sound waves interact with it. | TRUE | The strength of the echo (reflection) at a boundary is determined by the impedance mismatch. Large differences (e.g., between soft tissue and bone or air) cause strong reflections. | FALSE | The Doppler effect is the change in frequency of a wave due to motion. It is used to measure the velocity and direction of blood flow. | TRUE | Piezoelectric crystals within the transducer vibrate when a voltage is applied (creating sound) and generate a voltage when they are vibrated by returning echoes (receiving sound).
Pertaining to image display and archiving systems: | A "diagnostic workstation" used by radiologists typically has lower-resolution monitors than a clinical "review workstation". | The Window Width (WW) setting controls the overall brightness of a digital image. | PACS stands for "Picture Archiving and Communication System". | The DICOM file header contains only basic patient identification information. | A Look-Up Table (LUT) is a hardware device that stores the raw data from the scanner. | FALSE | A diagnostic workstation requires high-resolution, high-luminance, medical-grade monitors for primary interpretation. Review workstations may have lower specifications. | FALSE | The Window Level (WL) sets the midpoint of the displayed grayscale and controls the brightness. The Window Width (WW) controls the contrast. | TRUE | This is the correct term for the integrated system of computers, networks, and archives used to store, retrieve, distribute, and display medical images. | FALSE | The DICOM header contains extensive metadata, including patient demographics, study information, acquisition parameters, slice thickness, dose information, and more. | FALSE | A LUT is a set of instructions applied during image processing that maps the raw pixel values to specific brightness levels on the displayed image to modify contrast.
Regarding advanced imaging concepts and artifacts: | Time-of-Flight (TOF) is a PET technique that uses the Larmor frequency to improve image quality. | Beam hardening in CT causes dark streaks and bands to appear between two dense objects in an image. | Inversion Recovery (IR) is an MRI pulse sequence used to null the signal from a specific tissue type. | A "star artifact" in nuclear medicine is caused by patient motion during the scan. | Partial volume averaging can be reduced by using thicker reconstruction slices. | FALSE | Time-of-Flight PET uses the slight difference in the arrival time of the two 511 keV photons to better localize the annihilation event. Larmor frequency is an MRI concept. | TRUE | This artifact occurs because the beam's average energy increases as it passes through the first dense object, causing less-than-expected attenuation in the second object. | TRUE | By applying a 180-degree inversion pulse and waiting a specific time (TI), the longitudinal magnetization of a target tissue (like fat or fluid) can be nulled just before image acquisition. | FALSE | A "star artifact" is a form of reconstruction artifact that appears as streaks radiating from a very "hot" area (a source of high activity) in a SPECT or PET image. | FALSE | Partial volume averaging occurs when multiple tissue types occupy a single voxel. It is reduced by using thinner slices, which create smaller voxels.
Concerning radiation dose units and their application: | The Gray (Gy) is the unit used to quantify effective dose. | Equivalent dose (Sv) is calculated to account for the different radiosensitivities of various organs. | The radiation weighting factor (wR​) for protons is the same as for X-rays. | The Dose-Length Product (DLP) in CT is a measure of the total energy imparted over the entire scan volume. | A patient's occupational radiation dose limit includes dose received from personal medical examinations. | FALSE | The Gray (Gy) is the unit of absorbed dose. The Sievert (Sv) is the unit used for effective dose and equivalent dose. | FALSE | Equivalent dose accounts for the different biological effectiveness of various types of radiation (using wR​). Effective dose accounts for organ radiosensitivity (using wT​). | FALSE | The wR​ for X-rays is 1. Protons are more damaging and have a wR​ of 2 (formerly 5). Alpha particles have a wR​ of 20. | TRUE | DLP is calculated by multiplying the CTDIvol by the scan length. It is a good indicator of the total radiation imparted and is used to estimate effective dose (stochastic risk). | FALSE | Occupational dose limits apply only to dose received as part of a person's work. They explicitly exclude dose received from background radiation and personal medical procedures.
Regarding the physical basis of different imaging modalities: | CT imaging is based on the principle of nuclear magnetic resonance. | Ultrasound imaging relies on the transmission and reflection of mechanical waves. | SPECT imaging detects pairs of 511 keV photons emitted from positron-electron annihilation. | MRI contrast is primarily based on differences in tissue electron density. | Fluoroscopy provides real-time imaging using low-energy sound waves. | FALSE | CT imaging is based on measuring the attenuation of X-rays as they pass through the body from multiple angles. MRI is based on nuclear magnetic resonance. | TRUE | Ultrasound uses transducers to send pulses of high-frequency sound into the body and creates an image by detecting the echoes that reflect back from tissue interfaces. | FALSE | This describes PET imaging. SPECT (Single Photon Emission Computed Tomography) detects single gamma photons emitted directly from a radionuclide like Technetium-99m. | FALSE | MRI contrast is based on differences in tissue properties related to nuclear magnetic resonance, such as proton density, T1 and T2 relaxation times, flow, and diffusion. | FALSE | Fluoroscopy provides real-time imaging using a continuous or pulsed beam of X-rays, which are a form of ionizing electromagnetic radiation.
Concerning the process of X-ray production in a diagnostic tube: | Does it occur when moving electrons interact with target nuclei? | Is the process approximately 99% efficient in converting electron energy to X-rays? | Is the process more efficient with a rotating anode compared to a stationary anode? | Is the production rate increased if the atomic number of the target is increased? | Does it require a cooling air current at all times within the tube? | TRUE |  | FALSE | It is 99% inefficient; >99% of energy is converted to heat. | FALSE | A rotating anode is more efficient at dissipating heat, not at producing X-rays. | TRUE |  | FALSE | Cooling is required but not necessarily an air current; oil and other methods are used.
Which of the following factors will increase the radiation output from an X-ray tube? | Increasing the distance between the cathode and anode. | The addition of a filter to the beam. | Increasing the tube potential (kV). | Increasing the tube current (mA). | Using a constant potential generator compared to a single-phase waveform. | FALSE | No relation. | FALSE | Decreases total output. | TRUE | Output is approximately proportional to kV². | TRUE | Output is proportional to mA. | TRUE | kV is near its maximum for longer.
Regarding the internal components and thermal management of a diagnostic X-ray tube: | Is heat only removed to the tube envelope by the process of conduction? | Is the focusing cup negatively charged? | Are the rotor bearings lubricated with oil? | Is the anode stem designed to be a poor heat conductor? | Is rhenium added to the tungsten target to increase its toughness and lifespan? | FALSE | Heat is radiated through the vacuum to the envelope. | TRUE | To repel electrons from the filament and produce a narrow beam. | FALSE | Silver is used—oil would evaporate in the vacuum. | TRUE | Prevents damage to the rotor assembly. | TRUE | This alloy reduces surface pitting and increases lifespan.
Concerning the anode angle in a diagnostic X-ray tube: | Is it defined as the angle that the target face makes with the X-ray beam? | Is it the only factor that determines the size of the focal spot? | Is the angle generally in the range of 20-35 degrees? | Is the tube's heat rating increased if the angle is reduced? | Does it determine the size of the field covered by the X-ray beam at a given focus-to-film distance? | TRUE |  | FALSE | The size of the filament also determines focal spot size. | FALSE | The angle is generally 7-20°. | TRUE |  | TRUE | The steeper the angle, the narrower the x-ray beam and the smaller field covered.
With respect to the characteristics of an X-ray beam spectrum: | Is the spectrum unaffected by the use of filtration? | Does the spectrum vary with changes in tube current (mA)? | Is the maximum energy of the spectrum determined by the peak tube potential (kVp)? | Does the spectrum consist of both Bremsstrahlung and characteristic radiation if the kVp exceeds the target's K-edge energy? | Is the peak of the energy spectrum located at approximately two-thirds of the kVp? | FALSE | Filtration removes low energy photons, hardening the beam and increasing the mean energy of the spectrum. | FALSE | Increasing mA merely increases the number of photons. Their energy and therefore the shape of the spectrum remains the same if kV is unaltered. | TRUE | Peak potential gives the electrons striking the target their peak energy, therefore dictating peak photon energy in the x-ray beam. | TRUE | Bremsstrahlung is the continuous spectrum, and characteristic radiation the peaks specific to the anode material. | FALSE | The spectrum peak is ⅓ to ½ kVp, whereas average or effective energy is 50-60% of the maximum.
Considering the properties and function of an X-ray tube filament: | Must it have a high melting point and a low electrical resistance? | Do electrons evaporate from it through the process of thermionic emission? | Is tungsten used for the filament primarily because of its high atomic number? | Should it have a low vapour pressure to ensure a long lifespan? | Does it have a negative electrical potential relative to the anode? | FALSE | High melting point and high resistance to generate heat for thermionic emission. | TRUE | Filament heat allows electrons to boil off and create space charge around the cathode that can be accelerated towards the anode. | FALSE | This is the reason that it is chosen for the anode. | TRUE | A good property for thermionic emission. | TRUE | The focusing cup is usually more negative than the filament.
Regarding the principles and implications of the anode-heel effect: | Does it produce a completely uniform X-ray beam across its entire field? | Is the effect more prominent at the cathode end of the image receptor? | Is the effect more noticeable if the focus-to-film distance is increased? | Is the effect greater if the target angle is steeper (larger angle)? | Is it not used to any advantage in diagnostic radiography? | FALSE | Photons on the anode side of the beam have more target material to travel through, so are attenuated and the intensity reduced. | FALSE | By definition it is the 'anode' heel effect. | FALSE | With increased distance the beam diverges further and the film only intercepts the central part of the beam. | TRUE | The steeper the target, the further through the target material the photons on the anode side of the beam have to travel and the more attenuated they are. | FALSE | It is used in mammography and extremity radiography, where a thicker part of the body is placed towards the cathode side.
What are the effects of increasing the tube kV while keeping other factors constant? | It increases the patient's entrance surface dose. | It increases the ratio of scattered to primary radiation at the detector. | It increases the radiographic contrast. | It increases the degree of film blackening. | It increases the proportion of photoelectric interactions compared to Compton interactions. | TRUE |  | TRUE | Higher kV x-rays are more penetrating so scattering events occur deeper in the patient nearer the film. Also, the scattered photons are more penetrating. | FALSE | Contrast decreases as kV increases. | TRUE | Increased kV causes increased exposure and increased film density. | FALSE | At higher kVs Compton events are favoured.
With respect to the principles of X-ray beam attenuation: | Does increased tube filtration increase the half-value layer (HVL)? | Is the total attenuation the product of the Compton, photoelectric, and elastic attenuation effects? | Is the half-value thickness inversely proportional to the linear attenuation coefficient? | Is the attenuation altered when passing through materials of differing atomic numbers? | Is attenuation related to the inverse square law? | TRUE | Through beam hardening. | TRUE | This is the total attenuation. Attenuation coefficient is the sum of each process. | TRUE | Half value layer is 0.69/µ (µ = linear attenuation coefficient= fraction of the primary beam that is removed per unit distance). | TRUE | Attenuation is increased with increasing Z number and increasing density of the attenuating material, through increased Compton scatter and photoelectric absorption. (The probability of photoelectric absorption = z³). | FALSE | Attenuation is the reduction in intensity due to interactions in matter, whereas the inverse square law is the reduction in intensity due to beam divergence from a point source.
Regarding the concept of the Half-Value Layer (HVL): | Is it the thickness of a material that will reduce the intensity of a narrow X-ray beam to half of its original value? | Is it a measure of the penetrating power of an X-ray beam? | For a given X-ray energy, is the HVL of lead greater than that of aluminum? | Is the HVL reduced as the photon energy of the radiation decreases? | Does a narrow X-ray beam undergo exponential attenuation if it passes through successive half-value thicknesses? | FALSE | It is the thickness of a material that reduces intensity to half its original value. | TRUE |  | FALSE | The HVL decreases as the atomic number of the attenuating material increases. | TRUE |  | TRUE
Concerning the properties of the mass attenuation coefficient: | Is it measured in units of cm²/g? | Is it equal to the linear attenuation coefficient divided by the material's density? | Is it less for water than for ice? | Does it have many practical applications in diagnostic radiology? | Is it proportional to the linear attenuation coefficient? | TRUE |  | TRUE |  | FALSE | It is equal for water and ice, as it is independent of physical density. | FALSE | The linear attenuation coefficient has more practical applications, as film density produced by a certain depth of tissue is more useful than that produced by a certain mass of tissue. | TRUE
With respect to the linear attenuation coefficient (LAC): | Is it defined as the fractional reduction in beam intensity per unit thickness of material? | Can it be used to calculate the half-value thickness? | Does it increase as photon energy increases? | Is it measured in units of mm? | Is it true that a greater difference in the LAC between two tissues results in greater contrast between them? | TRUE |  | TRUE | HVT equals 0.693 divided by the linear attenuation coefficient. | FALSE | It decreases as photon energy increases. | FALSE | mm⁻¹ | TRUE | Contrast is proportional to the product of the difference between the 2 linear attenuation coefficients and the thickness of the tissues involved.
Considering the characteristics and effects of scattered radiation: | In diagnostic radiology, is more scatter measured on the X-ray tube side of the patient? | Is a Compton scattered photon deflected from its path with no loss of energy? | Does elastic scattering occur with no ionization of the atom? | Is a photoelectron produced during a Compton interaction? | At higher kV, are more photons deflected through large angles? | TRUE | Most interactions occur at the entrance surface of the patient and forward scattered photons are more attenuated than backscattered ones. | FALSE | This is elastic scattering. | TRUE |  | FALSE | These are formed in photoelectric interactions. | FALSE | Within the diagnostic range there is slight increase in forward scatter as kV increases, therefore there are more small angle deflections.
Regarding the mechanics of a Compton scattering event: | Does the interaction occur with a free or loosely bound electron? | Can the recoil electron be scattered in any direction, including backwards? | Is it true that the larger the angle of scatter, the greater the reduction in energy of the incident photon? | Can all of the photon's energy be transferred to the electron in this interaction? | Is the amount of scatter produced proportional to the electron density of the material? | TRUE | The incident photon interacts with an outer shell an electron. | FALSE | The scattered photon can be emitted in any direction, but the recoil electron can be projected only forwards or sideways. | TRUE |  | FALSE | The photon is scattered and therefore still retains some energy. Total absorption occurs in photoelectric interactions. | TRUE | The greater the concentration of electrons, the greater the probability of an interaction.
With respect to the photoelectric effect in the diagnostic energy range: | Does the incident photon completely disappear during the interaction? | Is it the main attenuation process in bone at 80 keV? | Can the interaction result in scattered radiation? | Does the probability of the interaction increase as the photon energy increases? | Does its contribution to the mass attenuation coefficient increase approximately as the cube of the atomic number? | TRUE |  | FALSE | The photoelectric effect predominates over Compton scatter in bone below 50kV and in soft tissue below 30kV. | FALSE | The photon is always completely absorbed. | FALSE | The probability of photoelectric interaction is inversely proportional to photon energy cubed. | TRUE | The probability of photoelectric interaction is proportional to Z³.
The photoelectric effect: | Does it involve interactions with free electrons? | Does it result in the production of characteristic radiation? | Does it only occur with electrons from the K shell? | Is it most important at the lower end of the diagnostic range of energies? | Does it result in the ionization of the atom? | FALSE | It involves inner shell electrons. | TRUE | The ejection of a photoelectron by the incident photon allows an outer shell electron to fall into the gap, with the emission of photons of characteristic radiation. | FALSE | Can also occur with the L shell. | TRUE |  | TRUE | As an electron is removed a net negative charge results.
Concerning the phenomenon of absorption edges: | For a given element, is the K-absorption edge energy greater than the L-absorption edge energy? | At photon energies just above the K-shell binding energy, is there a sharp decrease in the probability of photoelectric interactions? | Are K-absorption edges important when choosing an X-ray filter, a contrast medium, or an imaging phosphor? | Does an X-ray filter not transmit photons well if they have an energy equivalent to its own K-absorption edge? | Is the K-absorption edge of barium not relevant in diagnostic imaging because its atomic number is too high? | TRUE | K-shell binding energy is always greater than the L-shell binding energy for a given element. | FALSE | There is a sharp increase in the probability of photoelectric interactions just above the K-shell binding energy. | TRUE |  | FALSE | It will be relatively transparent to photons of the energy of its absorption edge. | FALSE | The atomic number of barium is 56 and the K-absorption edge is 37keV. Diagnostic x-ray beams contain a high proportion of photons around this energy, ensuring a high probability of photoelectric interactions.
Regarding the purpose and effects of X-ray beam filtration: | Does it reduce the maximum photon energy (kVp) of the beam? | Is filtration of the beam by the patient known as inherent filtration? | Does it improve the heat rating of the X-ray tube? | For filtering high-energy X-rays, is a copper filter more effective than an aluminum filter? | Does it result in an image with improved contrast? | FALSE | The kVp remains the same, but lower energy photons are filtered out and average kV increases. | FALSE | Inherent filtration results from absorption of x-rays as they pass through the tube. | FALSE | Tube loading is increased due to the reduced intensity of the x-ray beam once it has passed through the filter. | TRUE | Copper has a higher atomic number than aluminium, so is better at filtering higher energy x-rays. | FALSE | Filtration hardens the beam by increasing the mean energy of the photons, therefore contrast in the image is decreased.
With respect to the inherent filtration of an X-ray tube: | Does it primarily absorb high-energy X-rays? | Does it cause the beam to become "harder"? | Is it usually equivalent to 2.5 mm of aluminum? | Is the filtration increased if beryllium is used for the tube window instead of glass? | Is it mostly due to the insulating oil surrounding the tube? | FALSE | Absorbs low energy x-rays. | TRUE | By removing low energy photons and increasing the average energy of the beam. | FALSE | Total filtration is approximately 2.5mm Al, inherent is 0.5-1 mm. | FALSE | Beryllium has a lower atomic number than glass, therefore filtration is less. | FALSE | Due to target, tube window, and the oil.
Concerning the use and characteristics of added filtration: | Does it not affect the patient's entrance dose? | Does it alter the quality of the X-ray beam? | May it consist of a compound filter, such as copper and aluminum? | Is it generally made of aluminum in diagnostic X-ray tubes? | Does it not affect the intensity of the beam? | FALSE | Low energy photons are removed, which would contribute to patient dose, but not to the image. | TRUE | As mean kV is increased, penetrating power and therefore quality increases. | TRUE | Copper is used with a backing of aluminium on the patient side to absorb the 9kV characteristic radiation from the copper. | TRUE |  | FALSE | Intensity or amount of radiation is decreased by the filter.
Which of the following factors will increase the heat rating of an X-ray tube? | Using a rotating anode compared to a stationary anode. | Using a larger focal spot size. | Increasing the anode angle while keeping the effective focal spot size fixed. | Using half-wave rectification compared to full-wave rectification. | Producing heat more quickly. | TRUE | There is more efficient heat loss from a rotating anode, so the rating is higher. | TRUE | A larger focal spot causes less heating than if the beam were focused onto a smaller area. | FALSE | A smaller anode angle has a higher heat rating. | FALSE | Rating is increased with full wave rectification. | FALSE | This makes the tube rating lower.
Regarding methods for the measurement of radiation dose: | Can dose be read directly from photoconductive silicon diodes? | Can thermoluminescent dosimeters (TLDs) be used for both personal and patient dosimetry? | Can thimble ionization chambers be placed within the field of interest to measure dose? | Does a dose area product (DAP) meter provide a figure with the units of cGy·cm³? | Can staff dose be monitored using the photographic effect of silver bromide in a film badge? | TRUE | Useful for personal dosimeters and quality assurance. | TRUE |  | TRUE |  | FALSE | It is the product of dose and area with units cGy·cm². | TRUE
Considering the design and function of ionization chambers: | Are they used to measure absorbed dose in diagnostic radiography? | Does ionization of air in the chamber create a current by attracting ions to the chamber walls and central electrode? | Is air not an ideal material for the chamber because its effective atomic number differs greatly from that of soft tissue? | Do parallel plate ionization chambers, mounted on the collimator, measure the patient's dose area product? | Can any individual ionization chamber be used accurately over a very wide range of photon energies? | TRUE | They measure air kerma (kinetic energy released to matter), which at diagnostic energies is effectively equal to absorbed dose. | TRUE |  | FALSE | It is ideal, as the effective atomic number of air is 7.6, and that of soft tissue is 7.4. | TRUE |  | FALSE | Lower energy beams (such as in mammography) require ionization chambers with thinner walls to avoid undue attenuation of radiation by the chamber.
With respect to the use and properties of film badge dosimeters: | Do they use a double-emulsion film? | Are they calibrated using a cesium radiation source? | Do they have an open window for the detection of beta particles? | For staff monitoring, should they be analyzed once per year? | Do they use cadmium nuclei to detect neutron exposure? | TRUE | If the fast emulsion is over-exposed by a high-dose exposure, the slow emulsion can be read. | TRUE |  | TRUE |  | FALSE | They are subject to environmental effects, so should not be used for longer than a month. | TRUE | The interaction of neutrons with the cadmium nuclei results in gamma ray emission that is detected by the film.
Which of the following statements about thermoluminescent dosimeters (TLDs) are correct? | Is the phosphor commonly used lithium chloride? | Do X-ray interactions involve the outer shell (valence) electrons of the TLD's phosphor? | When exposed to radiation, do interactions excite electrons that become trapped in the forbidden energy band? | Does the amount of light produced when heated depend on the energy of the photons from the original exposure? | Is their dose response linear over a wide range? | FALSE | Lithium fluoride. | TRUE | Valence shell electrons are involved. | TRUE | A valence shell electron is excited into the conduction band and then falls back into an electron trap in the forbidden energy band. | TRUE |  | TRUE
Concerning the physical process of luminescence: | Is it the process where a material absorbs energy and re-emits it as light? | Is fluorescence defined as the delayed emission of light following energy input? | For light to be emitted from a phosphor, must electrons in electron traps fall to the conduction band? | After irradiation, must a thermoluminescent phosphor be stimulated with a laser for light to be emitted? | Is the intensity of light emitted from a phosphor proportional to the intensity of the irradiating X-ray beam? | TRUE |  | FALSE | Fluorescence is the instantaneous emission of light following energy input, phosphorescence is the delayed emission of light. | FALSE | The x-ray photons excite electrons from the valence band to the conduction band, where they then fall into electron traps. They must fall back to the valence band for light to be emitted. | FALSE | Thermoluminescence requires heating, photostimulable luminescence requires light. | TRUE
Which of the following are classified as deterministic effects of ionizing radiation? | Cataract formation | Epilation (hair loss) | Leukaemia | Lung cancer | Erythema (skin reddening) | TRUE | Cataract formation is a classic deterministic effect with a dose threshold to the lens of the eye. | TRUE | Epilation is a deterministic effect that occurs only after a threshold dose to the skin has been exceeded. | FALSE | Leukaemia is a primary example of a stochastic effect, where the probability of occurrence increases with dose, but there is no dose threshold. | FALSE | Like leukaemia, lung cancer is a stochastic effect of radiation, characterized by a probabilistic risk without a threshold dose. | TRUE | Erythema is a deterministic effect resulting from high radiation doses to the skin.
Which of the following are classified as stochastic effects of radiation? | Infertility | Leukaemia | Cataract formation | Cancer | Hair loss | FALSE | Infertility (sterility) is a deterministic effect, as it only occurs after a specific threshold dose to the gonads has been exceeded. | TRUE | Leukaemia is a radiation-induced cancer and is therefore a stochastic effect. | FALSE | Cataract formation is a well-known deterministic effect with a dose threshold to the lens of the eye. | TRUE | Radiation-induced cancer is the principal stochastic effect of concern in radiation protection. | FALSE | Hair loss (epilation) is a deterministic effect that occurs after a threshold dose to the skin has been reached.
Regarding the quantity Equivalent Dose: | Is it derived by multiplying the absorbed dose by a tissue weighting factor? | Is it measured in units of Sieverts (Sv)? | Is it calculated by averaging the dose over all tissues of the body? | Is it numerically the same as the absorbed dose for neutrons? | Does it account for the different sensitivities of tissues to radiation? | FALSE | It is derived from absorbed dose multiplied by a radiation weighting factor. | TRUE | The Sievert (Sv) is the dedicated SI unit for the quantities equivalent dose and effective dose. | FALSE | This describes effective dose. Equivalent dose is calculated for a specific organ or tissue. | FALSE | As the radiation weighting factor for neutrons ranges from 5 to 20, the equivalent dose will be 5 to 20 times greater than the absorbed dose. | FALSE | Effective dose takes into account the sensitivity of tissues to radiation.
Regarding the quantity Absorbed Dose: | Does the absorbed dose to an organ depend on the organ's mass? | Is it measured in units of Joules/kg? | Is it measured in units of Sieverts (Sv)? | Does its calculation depend on the radiation weighting factor? | Is it defined as the amount of energy deposited per unit mass in a medium? | FALSE | Absorbed dose is energy deposited per unit mass (J/kg). The total energy deposited depends on mass, but the dose itself is normalized by mass. | TRUE | The unit is Gray & 1 Gray= 1 Joule/Kg. | FALSE | The unit for absorbed dose is the Gray (Gy). | FALSE | This describes equivalent dose. Absorbed dose is a purely physical quantity and does not account for radiation type. | TRUE | This is the fundamental definition of absorbed dose.
Regarding the quantity Effective Dose: | Is it derived by multiplying the absorbed dose by a tissue weighting factor? | Is it measured in units of Gray (Gy)? | Does it take into consideration the different radiosensitivity of tissues? | Does it combine individual organ doses to provide a whole-body dose equivalent? | Is the effective dose from a dental film approximately 0.004 mSv? | FALSE | It is derived from equivalent dose multiplied by a tissue weighting factor. | FALSE | It is measured in Sieverts (Sv). | TRUE | This is its primary purpose, achieved by using tissue weighting factors for different organs. | TRUE | It sums the weighted equivalent doses in all irradiated tissues to provide a single value representing the overall stochastic risk. | TRUE | A single intraoral dental X-ray delivers a very small effective dose, typically in the range of 1-8 µSv (0.001-0.008 mSv).
Considering the properties of different types of ionizing radiation: | Are neutrons classified as low LET (linear energy transfer) radiation? | Is the radiation weighting factor for alpha particles equal to 20? | Do X-rays and beta particles have the same radiation weighting factor? | Is the radiation weighting factor for neutrons equal to 1 (unity)? | For X-rays, is the absorbed dose numerically equal to the equivalent dose? | FALSE | Neutrons are uncharged particles that cause dense ionization tracks via recoil protons, and are therefore high-LET radiation. | TRUE | Alpha particles are highly damaging, so they are assigned the highest radiation weighting factor (wR​) of 20. | TRUE | The radiation weighting factor for both is unity (1). | FALSE | It is 5-20 depending on the radiation energy. | TRUE | Equivalent dose = Absorbed dose x Radiation weighting factor, and the radiation weighting factor for x-rays is 1.
Which of the following tissues are assigned a high-risk tissue weighting factor (≥0.12) by the ICRP? | Colon | Skin | Breast | Bone marrow | Oesophagus | TRUE | The colon is assigned a tissue weighting factor (wT​) of 0.12. | FALSE | The skin has a low tissue weighting factor (wT​=0.01), reflecting a lower contribution to the overall stochastic risk compared to high-sensitivity organs. | TRUE | According to ICRP-103, breast is classified as high carcinogenic risk from radiation with a tissue-specific weighting factor of 0.12. | TRUE | Red bone marrow is assigned a tissue weighting factor (wT​) of 0.12. | FALSE | The oesophagus has a moderate tissue weighting factor (wT​=0.04) and is not considered in the highest risk category.
Which of the following tissues were assigned a moderate risk weighting factor (0.05) in older ICRP recommendations? | Skin | Gonads | Lung | Breast | Bone | FALSE | The skin is assigned a low tissue weighting factor (wT​=0.01). | FALSE | According to the recent amendment of ICRP June 2006 draft, gonads specific tissue weighting factor has decreased from 0.2 to 0.08. However, this is still classified as 'high risk'. | FALSE | The lung is considered a high-risk organ with a tissue weighting factor (wT​=0.12). | FALSE | The breast is considered a high-risk organ with a tissue weighting factor (wT​=0.12). | FALSE | Bone surfaces are assigned a low tissue weighting factor (wT​=0.01).
Are the following organ-specific tissue weighting factors correct according to modern ICRP recommendations? | Bone - 0.05 | Breast - 0.12 | Gonads - 0.12 | Thyroid - 0.05 | Skin - 0.01 | FALSE | It is 0.01. | TRUE | The breast is assigned a wT​ of 0.12, placing it in the highest sensitivity group. | FALSE | It is 0.08. | TRUE | The thyroid is assigned a wT​ of 0.05 (in older ICRP 60) or 0.04 (in newer ICRP 103). | TRUE | The skin is assigned a low sensitivity weighting factor of wT​=0.01.
Are the units for the following dosimetric quantities stated correctly? | Entrance Surface Dose (ESD) - Sv | Equivalent Dose - Gy | Dose Area Product (DAP) - Gy cm | Absorbed Dose - Joules/kg | Effective Dose - Gy | FALSE | The unit for ESD is Gray (Gy). | FALSE | The unit for Equivalent dose is Sieverts (Sv). | FALSE | The unit for DAP is Gy·cm². | TRUE | Absorbed dose is measured in Gray and 1Gy = 1J/kg. | FALSE | The unit for Effective dose is Sieverts (Sv).
Considering the characteristics of deterministic radiation effects: | Are diarrhoea and vomiting examples of deterministic effects? | Is there a threshold dose above which they do not occur? | Do these effects occur by chance? | Is the threshold dose the same for different deterministic effects? | Does the severity of the effect increase with increasing dose? | TRUE | These are symptoms of acute radiation syndrome, which is a deterministic effect. | FALSE | There is a threshold dose below which they do not occur. | FALSE | Stochastic effects occur by chance. Deterministic effects are a direct and expected consequence of a sufficiently high dose. | FALSE | The threshold dose varies for different deterministic effects (e.g., the threshold for skin erythema is different from that for cataracts). | TRUE | Once the threshold is exceeded, a higher dose results in a more severe outcome.
Considering the characteristics of stochastic radiation effects: | Is the probability of a stochastic effect independent of the radiation dose? | Do these effects occur immediately after exposure to ionizing radiation? | Are these effects described by a linear no-threshold (LNT) model? | Is sterility an example of a stochastic effect? | Is breast cancer an example of a stochastic effect? | FALSE | The probability increases with increasing dose. | FALSE | They occur after a latent period which lasts for many years. | TRUE | They occur by chance and their probability (but not severity) increases with the dose. The LNT model describes this relationship. | FALSE | Sterility is a deterministic effect because it requires a threshold dose to the gonads to occur. | TRUE | Radiation-induced cancer is the primary example of a stochastic effect.
Comparing the deterministic and stochastic effects of radiation: | In diagnostic radiology, are stochastic effects more common than deterministic effects? | Are the chances of producing deterministic effects the same for X-rays and gamma rays? | Is it true that no dose is considered completely safe for deterministic effects? | Can deterministic effects be non-additive over time? | Is skin necrosis an example of a stochastic effect? | FALSE | In diagnostic radiology, deterministic effects are virtually never seen as doses are kept far below thresholds. The risk of stochastic effects is the primary concern. | TRUE | Since the radiation weighting factor for both is one, they have the same biological effectiveness for producing deterministic effects. | FALSE | This is true for stochastic effects. Deterministic effects have a threshold, and any dose below that threshold is considered safe from that specific effect. | TRUE | Unlike stochastic risk, which is considered cumulative, some deterministic effects may be partially repaired by the body, so the damage is not strictly additive over long time periods. | FALSE | Skin necrosis (tissue death) is a severe deterministic effect that occurs at very high radiation doses.
Considering the physical properties of different types of ionizing radiation: | Do beta particles travel through matter at high speeds? | Do alpha particles travel through matter at low speeds? | Are alpha particles similar in structure to the nucleus of a hydrogen atom? | Are beta particles heavier than alpha particles? | Do alpha particles have useful applications in diagnostic radiology? | TRUE | Beta particles are high-energy electrons and travel at relativistic speeds. | TRUE | Alpha particles have a large mass and double charge making them travel slowly through matter. | FALSE | They are similar to a helium nucleus with 2 protons and 2 neutrons. A hydrogen nucleus is a single proton. | FALSE | Alpha particles are thousands of times heavier than beta particles (electrons). | FALSE | They produce a large amount of ionization per unit length of the medium through which they travel making them unsafe for use in radiology.
Are the following values representative of typical entrance surface doses (ESDs) for these examinations? | PA chest film - 0.15 mGy | Lateral lumbar spine X-ray - 12 mGy | AP skull X-ray - 2 mGy | AP abdomen film - 8 mGy | Fluoroscopy dose rate - 100-150 mGy/min | TRUE | This is a typical value for a well-calibrated adult PA chest X-ray. | TRUE | Due to the thickness of the anatomy in the lateral projection, a significantly higher exposure and entrance surface dose are required. | TRUE | The DRL is 3mGy, so 2 mGy is a typical value. | FALSE | Entrance surface dose is usually 5mGy. The DRL for an abdominal film is 7 mGy. | FALSE | The skin dose rate of fluoroscopy is 5-50 mGy/min.